## Sheet1
TRADE_NAME | FIRM_NAME | MANUFACTURER_RECALL_REASON | ACTION | Recall Type | processed_RECALL | processed_ACTION
Synapse PACS | Fujifilm Medical Systems U.S.A., Inc. | The software does not update measurements and calculations in the Clinical Reporting Application (CRA) when the ventricular trace is changed in the study by a different user. | Action On 12/11/2020 Fujifilm sent consignee notifications via FedEx Priority Overnight stating that there are issues with Synapse PACS Software Version 7.0.000 and 7.0.100. Fujifilm provided a work around until a hotfix is scheduled and completed. | SoftwareChangeControl | • The software's failure to update measurements and calculations in the Clinical Reporting Application (CRA) when the ventricular trace is changed by a different user leads to inaccurate or outdated actions. • Inconsistent actions are performed based on incorrect or outdated measurements and calculations, potentially affecting the quality and reliability of clinical reports. • The software's inability to reflect changes made by different users in real-time hinders collaborative actions and may lead to errors or discrepancies in the reporting process. • Users are not notified of changes made by others, causing potential confusion and misinformed actions based on outdated information. | • Take actions to work around issues with Synapse PACS Software Version 7.0.000 and 7.0.100 as provided by Fujifilm until a hotfix is scheduled and completed.
Synapse PACS | Fujifilm Medical Systems U.S.A., Inc. | FUJIFILM has become aware of the possibility that certain CT studies may report a much higher 3D Sphere MAX value than expected. This issue is specific to 3D Sphere - other density tools are not affected.****** | Action Letter was sent to all affected customers on 09/11/2020 stating the following: FUJIFILM will offer an upgrade to Synapse PACS 7.1.100 to resolve this issue. Synapse PACS 7.1.100 is estimated to be available by the end of September. ************Until the patch is available, users can continue to safely use Synapse PACS by following these additional instructions:************1. Ensure that the 3D Sphere MAX value is within the sphere***s range. If the 3D Sphere MAX value is observed to be much higher than expected, the value should not be used when rendering a clinical decision until the issue is corrected. ************Customers should complete and return the Field Action Verification Form attached to the recall letter. | SoftwareChangeControl | • CT studies may report incorrect 3D Sphere MAX values, leading to inaccurate results and actions. • Incorrect values may affect diagnosis and treatment decisions, leading to potential harm to patients and actions. • This issue may result in compromised confidence in test results, requiring re-testing and actions. | • To ensure safe use of Synapse PACS, take the following actions: ensure the 3D Sphere MAX value is within the sphere’s range and do not use the value when rendering a clinical decision if it is higher than expected until the issue is corrected. • Complete and return the Field Action Verification Form as part of these actions.
Securitas | Securitas Healthcare LLC | This is a risk that alarms will not be received by the Arial Wireless Nurse Call System and premature depletion of the battery. | Action Securitas Healthcare issued a recall notice on 06/19/2024 to its consignees via USPS certified mail. The notice explained the problem, potential hazard, and requested the following: ***"Customer Required Actions:***1. Closely Monitor Missing Device Alerts within the Arial Application and Immediately Verify Status of Missing Devices and Cease Using any Call Stations That Do Not Reconnect.***Clients should closely monitor the "Missing Device Alerts within the Arial Application and immediately check the status of Call Stations that trigger a Missing Device Alert and replace the Call Station battery if it is low or fully depleted.***Instructions for replacing the Call Station battery are attached hereto as Appendix 1. Instructions for reconnecting missing devices and confirming that the Call Stations are properly transmitting information to the Arial Emergency and Nurse Call System are attached hereto as Appendix 2. Clients should immediately cease using any Call Station that will not reconnect to the Arial Emergency and Nurse Call System.***2. Closely Monitor the Battery Status of all Call Stations in the Arial Web Client and Replace Battery if it is Low . Clients should closely monitor the Call Station Battery Status in the Arial Web Client and immediately replace the battery if it is low .***3. Consider Use of Supplementary Communication Devices or Procedures for High Risk Residents. Clients should consider enacting backup procedures, including increased room checks, or providing alternate emergency and nurse call communication means, such as a wearable pendant, for high-risk residents." ***"If you have distributed the product further, immediately notify your accounts that received the product identified above of this notification and ask them to contact Securitas Technical Support at 1-800-824-2996 (8:00 AM 5:00 PM CT, M-F) to obtain a response form and coordinate correction of affected devices." ***Securitas Healthcare is developing a firmware patch to remediate t | SoftwareDeployment | • Failure of alarms to be received by the Arial Wireless Nurse Call System may hinder prompt actions. • Premature depletion of the battery may lead to system downtime and inadequate actions. | • Customer Required Actions: Closely monitor Missing Device Alerts within the Arial Application and immediately verify the status of Missing Devices and cease using any Call Stations that do not reconnect. • Customer Required Actions: Closely monitor the battery status of all Call Stations in the Arial Web Client and replace the battery if it is low. • Customer Required Actions: Consider enacting backup procedures, including increased room checks, or providing alternate emergency and nurse call communication means, such as a wearable pendant, for high-risk residents. • Customer Required Actions: If you have distributed the product further, immediately notify your accounts that received the product and ask them to contact Securitas Technical Support to obtain a response form and coordinate correction of affected devices.
Securitas | Securitas Healthcare LLC | There is a risk that alarms will not be received by the Arial Wireless Emergency Call System and premature depletion of the battery. | Action Securitas Healthcare issued a recall notice on 06/19/2024 to its consignees via USPS certified mail. The notice explained the problem, potential hazard, and requested the following: ***"Customer Required Actions:***1. Reduce the Missing Device Timeout to 125 Minutes Hours for all Call Stations.***Clients should reduce the Missing Device Timeout to 125 minutes so that a Missing Device Alert will be triggered if a Call Station fails to send a supervision message to the Arial Emergency and Nurse Call System within 125 minutes from the last supervision message. Instructions for reducing the Missing Device Timeout to 125 minutes are attached hereto as Appendix 1.***2. Closely Monitor Missing Device Alerts within the Arial Web Application and Immediately Verify Status of Missing Devices and Cease Using any Call Stations That Do Not Reconnect.*** Clients should closely monitor the "Missing Device Alerts within the Arial Web Application and immediately check the status of Call Stations that trigger a Missing Device Alert and replace the Call Station battery if it is low or fully depleted.***Instructions for replacing the Call Station battery are attached hereto as Appendix 2. Instructions for reconnecting missing devices and confirming that the Call Stations are properly transmitting information to the Arial Emergency and Nurse Call System are attached hereto as Appendix 3. Clients should immediately cease using any Call Station that will not reconnect to the Arial Emergency and Nurse Call System.***3. Closely Monitor the Battery Status of all Call Stations in the Arial Web Client and Replace Battery if it is Low .***Clients should closely monitor the Call Station Battery Status in the Arial Web Client and immediately replace the battery if it low as described in the documentation.***4. Use Supplementary Communication Devices or Procedures for High Risk Residents.***Clients should enact backup procedures, including increased room checks, or providing alternate emergency | SoftwareDeployment | • Failure of alarms to be received by the Arial Wireless Emergency Call System requires urgent actions. • Premature depletion of the battery affects the system's performance and user safety, necessitating prompt actions. | • Customer Required Actions: Reduce the Missing Device Timeout to 125 minutes for all Call Stations to trigger a Missing Device Alert if a Call Station fails to send a supervision message. • Customer actions: Closely monitor Missing Device Alerts and immediately verify the status of Missing Devices, ceasing use of any Call Stations that do not reconnect. • Customer actions: Closely monitor the Battery Status of all Call Stations and replace the battery if it is low. • Customer actions: Enact supplementary communication devices or procedures for High-Risk Residents, including increased room checks or providing alternate emergency procedures.
VITROS 4600 Immunodiagnostic System | Ortho-Clinical Diagnostics, Inc. | A software defect affecting VITROS Systems running VITROS Software Versions 3.8.0 and 3.8.1 causes the Configurable Westgard Rules to not report as expected. When a new control is defined, configurable Westgard rules, if selected, are not flagged or reported with the applicable QC results if the statistical criteria is met. This may result in acceptance of QC results that do not meet acceptance criteria, which may cause erroneous patient results to be reported. | Action On January 26, 2024, the firm issued a customer letter to all customers who have V3.8.0 or V3.8.1 installed on their VITROS Analyzer. Customers were informed of the issue and instructed to follow the instructions when using the VITROS System onboard QC package to use the configurable Westgard Rules.******The software will be updated to mitigate this issue within the VITROS Software QC feature. The estimated availability is 2Q2024.******If you have further questions, please contact Global Services Organization at 1-800-421-3311. | SoftwareDesignChange | • Configurable Westgard Rules do not report as expected, resulting in unflagged or unreported QC results that meet statistical criteria, allowing acceptance of unacceptable results and potentially causing erroneous patient results through actions. • Failure to flag or report applicable QC results if statistical criteria is met, leading to acceptance of QC results that do not meet acceptance criteria, which can cause erroneous patient results to be reported through actions. | • Follow the instructions when using the VITROS System onboard QC package to use the configurable Westgard Rules as part of your actions. • Take the actions to mitigate this issue within the VITROS Software QC feature. • If you have further questions, take actions to contact Global Services Organization as needed.
VITROS XT 7600 Integrated System | Ortho-Clinical Diagnostics, Inc. | A software defect affecting VITROS Systems running VITROS Software Versions 3.8.1 causes QC baseline statistics to not update as expected when changed by the user. This may cause erroneous patient results to be reported. | Action On January 26, 2024, the firm sent a customer letter to all customers who have V3.8.1 installed on their VITROS Analyzer. Customers were informed of this issue and instructed to follow the instructions when using the VITROS System onboard QC package to update the QC baseline.******The software will be updated to mitigate both issues within the VITROS Software QC feature. The estimated availability is 2Q2024.******If you have further questions, please contact Global Services Organization at 1-800-421-3311. | SoftwareDesignChange | • Erroneous patient results may be reported due to QC baseline statistics not updating as expected after user actions. • Unexpected behavior of VITROS Systems may lead to inaccurate results and compromised system performance following user actions. | • Follow the instructions when using the VITROS System onboard QC package to update the QC baseline through actions. • Take actions to update the QC baseline according to the instructions provided. • Take further actions by contacting Global Services Organization at 1-800-421-3311 if you have any questions.
VITROS XT 3400 Chemistry System | Ortho-Clinical Diagnostics, Inc. | A software defect affecting VITROS Systems running VITROS Software Versions 3.8.0 and 3.8.1 causes the Configurable Westgard Rules to not report as expected. When a new control is defined, configurable Westgard rules, if selected, are not flagged or reported with the applicable QC results if the statistical criteria is met. This may result in acceptance of QC results that do not meet acceptance criteria, which may cause erroneous patient results to be reported. | Action On January 26, 2024, the firm issued a customer letter to all customers who have V3.8.0 or V3.8.1 installed on their VITROS Analyzer. Customers were informed of the issue and instructed to follow the instructions when using the VITROS System onboard QC package to use the configurable Westgard Rules.******The software will be updated to mitigate this issue within the VITROS Software QC feature. The estimated availability is 2Q2024.******If you have further questions, please contact Global Services Organization at 1-800-421-3311. | SoftwareDesignChange | • Configurable Westgard rules do not report as expected, causing QC results that do not meet acceptance criteria to be accepted, which may lead to erroneous patient results and incorrect actions. • Failure to flag or report applicable QC results when statistical criteria is met, resulting in inaccurate conclusions and misguided actions. | • Follow the instructions when using the VITROS System onboard QC package to use the configurable Westgard Rules, which are the required actions. • Take the actions of contacting the Global Services Organization at 1-800-421-3311 if you have further questions.
VITROS XT 7600 Integrated System | Ortho-Clinical Diagnostics, Inc. | A software defect affecting VITROS Systems running VITROS Software Versions 3.8.0 and 3.8.1 causes the Configurable Westgard Rules to not report as expected. When a new control is defined, configurable Westgard rules, if selected, are not flagged or reported with the applicable QC results if the statistical criteria is met. This may result in acceptance of QC results that do not meet acceptance criteria, which may cause erroneous patient results to be reported. | Action On January 26, 2024, the firm issued a customer letter to all customers who have V3.8.0 or V3.8.1 installed on their VITROS Analyzer. Customers were informed of the issue and instructed to follow the instructions when using the VITROS System onboard QC package to use the configurable Westgard Rules.******The software will be updated to mitigate this issue within the VITROS Software QC feature. The estimated availability is 2Q2024.******If you have further questions, please contact Global Services Organization at 1-800-421-3311. | SoftwareDesignChange | • Failure to flag or report configurable Westgard rules actions may lead to acceptance of QC results that do not meet acceptance criteria. • Erroneous patient results actions may be reported due to the defective software. | • Follow the instructions when using the VITROS System onboard QC package to use the configurable Westgard Rules for actions. • Contact Global Services Organization at 1-800-421-3311 for further actions.
VITROS 3600 Immunodiagnostic System | Ortho-Clinical Diagnostics, Inc. | A software defect affecting VITROS Systems running VITROS Software Versions 3.8.1 causes QC baseline statistics to not update as expected when changed by the user. This may cause erroneous patient results to be reported. | Action On January 26, 2024, the firm sent a customer letter to all customers who have V3.8.1 installed on their VITROS Analyzer. Customers were informed of this issue and instructed to follow the instructions when using the VITROS System onboard QC package to update the QC baseline.******The software will be updated to mitigate both issues within the VITROS Software QC feature. The estimated availability is 2Q2024.******If you have further questions, please contact Global Services Organization at 1-800-421-3311. | SoftwareDesignChange | • QC baseline statistics do not update as expected when changed by the user, leading to actions that cause inaccurate patient results. • Erroneous patient results are reported due to actions not reflecting updated QC baseline statistics, posing a risk to patient safety. | • Follow the instructions when using the VITROS System onboard QC package to update the QC baseline, which includes actions to mitigate the issue. • Take actions to update the VITROS Software QC feature when the updated software becomes available. • Contact the Global Services Organization at 1-800-421-3311 if you have further questions or require assistance with the actions.
VITROS 3600 Immunodiagnostic System | Ortho-Clinical Diagnostics, Inc. | A software defect affecting VITROS Systems running VITROS Software Versions 3.8.0 and 3.8.1 causes the Configurable Westgard Rules to not report as expected. When a new control is defined, configurable Westgard rules, if selected, are not flagged or reported with the applicable QC results if the statistical criteria is met. This may result in acceptance of QC results that do not meet acceptance criteria, which may cause erroneous patient results to be reported. | Action On January 26, 2024, the firm issued a customer letter to all customers who have V3.8.0 or V3.8.1 installed on their VITROS Analyzer. Customers were informed of the issue and instructed to follow the instructions when using the VITROS System onboard QC package to use the configurable Westgard Rules.******The software will be updated to mitigate this issue within the VITROS Software QC feature. The estimated availability is 2Q2024.******If you have further questions, please contact Global Services Organization at 1-800-421-3311. | SoftwareDesignChange | • Configurable Westgard rules are not flagged or reported with applicable QC results when statistical criteria are met, leading to inadequate quality control actions. • Erroneous patient results may be reported due to acceptance of QC results that do not meet acceptance criteria, resulting in inaccurate medical actions. • The defect may cause the system to malfunction, leading to unreliable test results and inadequate corrective actions. | • Take actions to follow the instructions when using the VITROS System onboard QC package to use the configurable Westgard Rules. • Take actions to contact Global Services Organization at 1-800-421-3311 if you have further questions.
VITROS XT 3400 Chemistry System | Ortho-Clinical Diagnostics, Inc. | A software defect affecting VITROS Systems running VITROS Software Versions 3.8.1 causes QC baseline statistics to not update as expected when changed by the user. This may cause erroneous patient results to be reported. | Action On January 26, 2024, the firm sent a customer letter to all customers who have V3.8.1 installed on their VITROS Analyzer. Customers were informed of this issue and instructed to follow the instructions when using the VITROS System onboard QC package to update the QC baseline.******The software will be updated to mitigate both issues within the VITROS Software QC feature. The estimated availability is 2Q2024.******If you have further questions, please contact Global Services Organization at 1-800-421-3311. | SoftwareDesignChange | • QC baseline statistics do not update as expected when changed by the user's actions. • Erroneous patient results may be reported due to the defective software's actions. | • Follow the instructions when using the VITROS System onboard QC package to update the QC baseline, which includes actions to mitigate issues within the VITROS Software QC feature. • Take actions to update the software to resolve the issues within the VITROS Software QC feature. • If you have further questions, contact Global Services Organization at 1-800-421-3311 for actions regarding the issue.
VITROS 4600 Immunodiagnostic System | Ortho-Clinical Diagnostics, Inc. | A software defect affecting VITROS Systems running VITROS Software Versions 3.8.1 causes QC baseline statistics to not update as expected when changed by the user. This may cause erroneous patient results to be reported. | Action On January 26, 2024, the firm sent a customer letter to all customers who have V3.8.1 installed on their VITROS Analyzer. Customers were informed of this issue and instructed to follow the instructions when using the VITROS System onboard QC package to update the QC baseline.******The software will be updated to mitigate both issues within the VITROS Software QC feature. The estimated availability is 2Q2024.******If you have further questions, please contact Global Services Organization at 1-800-421-3311. | SoftwareDesignChange | • Erroneous patient results are reported due to QC baseline statistics not updating as expected when changed by the user's actions. • Software defect in VITROS Systems running VITROS Software Versions 3.8.1 affects system performance, causing inaccurate results, which may impact patient safety through incorrect actions. | • Follow the instructions when using the VITROS System onboard QC package to update the QC baseline actions.
VITROS 5600 Immunodiagnostic System | Ortho-Clinical Diagnostics, Inc. | A software defect affecting VITROS Systems running VITROS Software Versions 3.8.0 and 3.8.1 causes the Configurable Westgard Rules to not report as expected. When a new control is defined, configurable Westgard rules, if selected, are not flagged or reported with the applicable QC results if the statistical criteria is met. This may result in acceptance of QC results that do not meet acceptance criteria, which may cause erroneous patient results to be reported. | Action On January 26, 2024, the firm issued a customer letter to all customers who have V3.8.0 or V3.8.1 installed on their VITROS Analyzer. Customers were informed of the issue and instructed to follow the instructions when using the VITROS System onboard QC package to use the configurable Westgard Rules.******The software will be updated to mitigate this issue within the VITROS Software QC feature. The estimated availability is 2Q2024.******If you have further questions, please contact Global Services Organization at 1-800-421-3311. | SoftwareDesignChange | • Configurable Westgard rules do not report as expected, causing incorrect acceptance of QC results that do not meet acceptance criteria, which leads to erroneous patient results through unacceptable actions. • Failure to flag or report configurable Westgard rules when selected and statistical criteria are met results in inadequate quality control, risking inaccurate patient diagnosis through unacceptable actions. | • Follow the instructions when using the VITROS System onboard QC package to use the configurable Westgard Rules for immediate actions. • Contact Global Services Organization at 1-800-421-3311 for further actions.
VITROS 5600 Immunodiagnostic System | Ortho-Clinical Diagnostics, Inc. | A software defect affecting VITROS Systems running VITROS Software Versions 3.8.1 causes QC baseline statistics to not update as expected when changed by the user. This may cause erroneous patient results to be reported. | Action On January 26, 2024, the firm sent a customer letter to all customers who have V3.8.1 installed on their VITROS Analyzer. Customers were informed of this issue and instructed to follow the instructions when using the VITROS System onboard QC package to update the QC baseline.******The software will be updated to mitigate both issues within the VITROS Software QC feature. The estimated availability is 2Q2024.******If you have further questions, please contact Global Services Organization at 1-800-421-3311. | SoftwareDesignChange | • The software defect prevents QC baseline statistics from updating as expected when changed by the user, leading to inaccurate actions. • Erroneous patient results may be reported due to the defect, compromising patient safety and requiring corrective actions. • The defect affects the system's performance, causing unreliable results and necessitating immediate actions. | • Follow the instructions when using the VITROS System onboard QC package to update the QC baseline for necessary actions. • Take actions to update the VITROS Software QC feature to mitigate issues. • If you have further questions, contact Global Services Organization at 1-800-421-3311 to take necessary actions.
smiths medical Medfusion | Smiths Medical ASD Inc. | Medfusion Model 4000 syringe infusion pumps, with the following software versions, may have the following issues: 1. Delivery During Motor Not Running High Priority Alarm (v1.0.0, v1.1.0, v1.1.1, v1.1.2), 2. Re-administered Loading Dose (v1.0.0, v1.1.0, v1.1.1, v1.1.2), 3. Incorrect Critical Data Failure Alarm, 4. Interruption of Bolus or Loading Dose Delivery (v1.0.0, v1.1.0, v1.1.1, v1.1.2), 5. Incorrect Total Bolus/Loading Dose Displayed (v1.1.0, v1.1.1, v1.1.2), 6. Volume Limit Before Bolus/Loading Dose Complete (v1.1.2), 7. Drug Library Lower Limit Displayed Incorrectly (v1.6.0, v1.6.1), 8. Depleted Battery Alarm (All versions previous to v1.6.5), 9. Loss of Wireless Connectivity (v1.5.0, v1.5.1, v1.6.0, v1.6.1, v1.6.4), 10. PharmGuard Server Password (v2.3, v2.4, v2.5). ***Smiths Medical corrected all issues included in this notification in previous software updates and the corrections were carried forward into all subsequent software releases.***Please ensure you have the most recent Medfusion software installed on your pumps.*** | Action Smiths Medical issued an URGENT MEDICAL DEVICE CORRRECTION notice to its consignees on 12/19/2023 by letter delivered via traceable means. This notification details the issues and the affected software versions. Please note that the pump displays the installed software version on the startup screen after the pump is powered on. The notice explained the issues with the pumps, the potential risks involved with their use, and requested the following actions be taken: ***1. Locate all affected pumps in your possession and ensure all users or potential users***of these devices are immediately made aware of this notification and proposed***mitigations. Please ensure you have the most recent Medfusion software installed on***your pumps. ***2. Complete and return the attached Response Form to smithsmedical8045@sedgwick.com within ten days of receipt to acknowledge your understanding of this notification. ***3. DISTRIBUTORS: If you have distributed potentially affected products to your customers, please immediately forward this notice to them. Request that they complete the response form and return it to smithsmedical8045@sedgwick.com. *** ***To report adverse events or product complaints: Global Complaint Management, globalcomplaints@icumed.com, 1-(866)-216-8806******Additional information or technical assistance: Technical Support, icumed.custhelp.com/app/market-action******Questions about this Field Correction Notice: Field Corrections, icumed.custhelp.com/app/market-action | SoftwareDesignChange | • Delivery actions interrupted due to High Priority Alarm. • Incorrect delivery actions of loading dose due to re-administration. • Incorrect actions triggering Critical Data Failure Alarm. • Interruption of bolus or loading dose delivery actions. • Incorrect display of total bolus/loading dose, affecting delivery actions. • Volume limit actions restricted before bolus/loading dose completion. • Incorrect display of drug library lower limit, affecting dosing actions. • Depleted battery actions causing alarm and system failure. • Loss of wireless connectivity actions affecting remote monitoring and control. • PharmGuard Server password actions compromised, affecting security. | • Locate all affected pumps in your possession and ensure all users or potential users are immediately made aware of this notification and proposed mitigations as part of the required actions. • Complete and return the attached Response Form to acknowledge your understanding of this notification as part of the actions. • If you have distributed potentially affected products to your customers, forward this notice to them and request that they complete the response form and return it as part of the actions.
smiths medical Medfusion | Smiths Medical ASD Inc. | Medfusion syringe pumps, with software versions before v6.0.0, may have the following issues:1. Delivery During Motor Not Running High Priority Alarm, 2. Infusion Restarted with Incorrect Parameters, 3. Screen Lock, 4. Interruption of Bolus or Loading Dose Delivery, 5. Pump Displays Incorrect Bolus/Loading Dose, 6. Loading/Bolus Dose Below the Minimum Recommended Rate, 7. Motor Rate Error, 8. Incorrect Recall Last Settings, 9. Corrupt Configuration, 10. Auto Lock, 11. Toolbox Configuration Loading Dose Time Values.******Smiths Medical corrected all issues included in this notification in previous software updates and the corrections were carried forward into all subsequent software releases. Please ensure you have the most recent Medfusion software (v6.0.0) installed on your pumps. | Action Smiths Medical issued an URGENT MEDICAL DEVICE CORRRECTION notice to its consignees on 12/16/2023 by letter delivered via traceable means. The notice explained the issues with the pumps, the potential risks involved with their use, and requested the following actions be taken: ***1. Locate all affected pumps in your possession and ensure all users or potential users***of these devices are immediately made aware of this notification and proposed***mitigations. Please ensure you have the most recent Medfusion software installed on***your pumps.***2. Complete and return the attached Response Form to smithsmedical8045@sedgwick.com***within ten days of receipt to acknowledge your understanding of this notification.***3. DISTRIBUTORS: If you have distributed potentially affected products to your customers,***please immediately forward this notice to them. Request that they complete the response form and return it to smithsmedical8045@sedgwick.com. ***Inquires: ***To report adverse events or product complaints: Global Complaint Management, globalcomplaints@icumed.com, 1-(866)-216-8806******Additional information or technical assistance: Technical Support, icumed.custhelp.com/app/market-action******Questions about this Field Correction Notice: Field Corrections, icumed.custhelp.com/app/market-action | SoftwareDesignChange | • Failure of Medfusion syringe pumps to deliver medication accurately and safely due to software issues, affecting patient treatment actions. • Incorrect infusion settings and parameters, leading to potential harm or injury to patients during treatment actions. • Inability of pumps to display correct bolus/loading dose information, causing potential medication errors and affecting patient safety actions. • Motor rate errors and interruptions to bolus or loading dose delivery, impacting the effectiveness of treatment actions. | • Take the following actions: Locate all affected pumps in your possession and ensure all users or potential users are immediately made aware of this notification and proposed mitigations. • Take the following actions: Ensure you have the most recent Medfusion software installed on your pumps. • Take the following actions: Complete and return the attached Response Form to acknowledge your understanding of this notification. • Take the following actions: If you have distributed potentially affected products to your customers, please immediately forward this notice to them and request that they complete the response form and return it.
RayStation | RAYSEARCH LABORATORIES AB | RaySearch Laboratories AB became aware of an issue with the dose calculation for some imported VMAT plans where controls points are defined with wide gantry angle spacing. | Action RaySearch sent a Device Correction letter dated May 8, 2015 via e-mailed to their customers. The letter identified the affected product, problem, actions to be taken by users and contact information. | SoftwareDesignChange | • Inaccurate dose calculation occurs for certain VMAT plans due to wide gantry angle spacing, requiring corrective actions. • This issue affects the system's operation, potentially leading to inaccurate treatment plans and compromising patient safety, necessitating urgent corrective actions. | • Identify the affected product and the problem. • Take actions to correct the device as instructed by RaySearch. • Contact RaySearch for further information or assistance with actions to be taken.
RayStation | RAYSEARCH LABORATORIES AB | RaySearch Laboratories has recalled RaySearch Ray Station 4.0 software due to clinical dose calculation errors during radiation therapy. | Action On February 6, 2014, RaySearch Laboratories distributed Urgent Medical Device Correction notices to their customers via email regarding the affected product. To determine if the version you are using is affected, bring up the About box in the RayStation application and check if the number reported is "4.0.2.9". If so, the notice applies to your version. Customers are advised to inform all planning staff and all users about the instructions listed on the notice. Customers are advised to inspect their system and identify all installed units with the referenced software version numbers listed on the notice. Corrective labeling is distributed to the customers in the form of the safety notification attached to this report and the software problem will be resolved in RayStation patch 4.0.3 scheduled for market release end of February 2014 and in all future releases.****** For further information and return of the reply form, please contact: Freddie Cardel, Director of Customer Support, RaySearch Americas, (877) 778-3849 or freddie.cardel@raysearchlabs.com. | SoftwareDesignChange | • Clinical dose calculation errors occur during radiation therapy, leading to inaccurate actions. • Inaccurate dose calculations may result in incorrect treatment actions. | • Determine if your version is affected by checking the About box in the RayStation application for the number "4.0.2.9", and take actions accordingly. • Inform all planning staff and all users about the instructions listed on the Urgent Medical Device Correction notice, and take actions. • Inspect your system and identify all installed units with the referenced software version numbers listed on the notice, and take actions. • Implement the corrective labeling distributed in the safety notification, and take actions. • Contact Freddie Cardel, Director of Customer Support, RaySearch Americas, for further information and to return the reply form, and take actions.
RayStation | RAYSEARCH LABORATORIES AB | Potential for dose errors due to software program errors. | Action RaySearch Laboratories sent an Field Safety Notice, Medical Device Correction letter dated February 4, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.************ACTIONS TO BE TAKEN BY THE USER******A workaround that eliminates the described error is to select the ***discard rotations*** option in the Image Registration Module before creating a rigid registration intended to be used for hybrid deformable registration. Another option is to use the ROI based deformable registration when applicable, as this functionality is not affected by the error.************SOLUTION******This issue will be resolved in a patch release of RayStation, scheduled for market release February 28, 2014. Release schedule is subject to regulatory approval for some markets. In the meantime, this field safety notice is being distributed to all customers.************CONTACT INFORMATION******For further information, please contact******For the Americas:******Customer Support, RaySearch Americas, at +1 877 778 3849 ************For Europe, Asia and rest of the world:******RaySearch Laboratories AB, at +46 8 5450 6130 ************Thank you for your cooperation, and we apologize for any inconvenience | SoftwareDesignChange | • Software program errors may lead to incorrect dosing, requiring corrective actions. • Dose errors can result in patient harm or adverse health effects, necessitating immediate actions. • Software malfunction can compromise the system's performance, demanding prompt actions. | • To eliminate the described error, take the following actions: select the discard rotations option in the Image Registration Module before creating a rigid registration intended to be used for hybrid deformable registration. • As an alternative, take the actions of using the ROI based deformable registration when applicable, as this functionality is not affected by the error.
Merge PACS | Merge Healthcare, Inc. | When cut lines are shown on an image, the lines may go from being vertical and then shift to horizontal causing confusion to the physician. This may cause a misdiagnosis. | Action Merge sent an Urgent Medical Device Recall letter dated May 5, 2017, to all affected customers. Customers were instructed to be aware of the following workaround: Do not have the simple ***close study*** red ***x icon available in the reading physician***s toolbars. This will force them to set status on close which also presents the warning ***Not all images have been reviewed, do you want to set status anyway and close study or continue reviewing unviewed images?*********YOUR RESPONSE TO THIS NOTIFICATION IS REQUIRED******By replying to this notification, you will be taking the necessary steps to ensure that the issue listed in this document will be resolved.******Please reply using the enclosed form and the return addressed envelope.******Customers with questions were instructed to send an email to recall@merge.com.******For questions regarding this recall call 877-741-6369. | SoftwareDesignManufacturingProcess | • Cut lines on an image can shift from vertical to horizontal, causing confusion to the physician during diagnosis, leading to incorrect actions. • The misalignment of cut lines can result in misdiagnosis, leading to incorrect actions. | • To take necessary actions, do not have the simple ***close study*** red ***x icon available in the reading physician***s toolbars, which will force you to set status on close and present a warning. • Take actions by replying to this notification using the enclosed form and return addressed envelope to ensure the issue is resolved. • For questions, take actions by sending an email to recall@merge.com. • Take actions by calling 877-741-6369 for questions regarding this recall.
Merge PACS | Merge Healthcare, Inc. | Merge PACS did not show unviewed images when the last view was skipped. There is a potential risk to health of a physician misdiagnosis because not all images are available for viewing. | Action Merge Healthcare sent an Urgent Medical Device Recall letter dated May 5, 2017.************The letter explains the issue and informs the customer to be aware of the workaround. The customer is also notified a software fix for the issue was released in Merge PACS 7.1.1 (or later). The customer is requested to ensure all users of the product are provided with the notification and if it has been further distributed, to notify downstream customers at once. The letter informs the customer a response is required 15 calendar days after receipt of the recall letter. The enclosed response asks, in part, if the customer is interested in accepting the fix and if not, to state the reason why. ************The recalling firm will contact all unresponsive customers a second and third time. Methods for contacting the customer may be by phone or email. ************For further questions, please call (877) 741-5369. | SoftwareDesignManufacturingProcess | • Merge PACS failure to display unviewed images when the last view is skipped leads to incomplete information for physicians, potentially resulting in misdiagnosis actions. • Incomplete image availability affects the accuracy of medical diagnoses, posing a risk to patients' health and well-being actions. | • Be aware of the workaround and take actions to ensure all users of the product are provided with the notification. • If the product has been further distributed, notify downstream customers at once and take actions to inform them. • Respond to the recall letter and take actions to indicate whether you are interested in accepting the fix. • If not, take actions to state the reason why. • Take actions to contact the recalling firm for further questions.
MOSAIQ Oncology Information System | Elekta, Inc. | Image offsets calculated by oncology Information System *** software may be incorrect for Linear Accelerator (LINACs) not characterized at IEC61217 scale. | Action On January 15, 2025 Elekta issued a recall notification to affected consignees via Email. Elekta asked consignees to take the following actions. ******1. Customers included in the scope of the Field Safety Notice, should not calculate shifts from 2D verification images taken at a non-zero patient support angle (couch angle).******2. Please post this notice in a place accessible to all users, e.g. Instructions for Use, until this action is closed.******3. Advise the appropriate personnel, working with this product, on the content of this letter.******4. You are required to either acknowledge receipt of this notification via***the Elekta Care" Community or complete this form and return it to Elekta immediately upon receipt, but no later than within 30 days. | SoftwareDesign | • Inaccurate image offsets may lead to incorrect treatment planning actions. • LINACs not characterized at IEC61217 scale may cause unreliable calculations actions. • Incorrect calculations can result in compromised patient safety and treatment outcomes actions. | • Take the following actions: do not calculate shifts from 2D verification images taken at a non-zero patient support angle (couch angle). • Take the following actions: post this notice in a place accessible to all users, e.g. Instructions for Use, until this action is closed. • Take the following actions: advise the appropriate personnel, working with this product, on the content of this letter. • Take the following actions: acknowledge receipt of this notification via the Elekta Care" Community or complete this form and return it to Elekta immediately.
epoc Blood Analysis System | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics is recalling their epoc Blood Analysis System because of a software malfunction that occurs when specific criteria are met that can result in the incorrect patient name being assigned to a blood measurement by the device if the user of the device does not ensure verification of the correct patient name. The software malfunction has the potential to impact patient safety via misdiagnosis and/or mistreatment by a clinician if the wrong patient information is saved with the test result from the subject device. | Action Siemens issued Urgent Medical Device Correction (UMDC) letter to U.S. customers by FedEx beginning on November 22, 2024. Urgent Field Safety Notice (UFSN) was issued to OUS Siemens reporting sites with affected customers, for distribution and regional***reporting to their local competent authorities beginning on November 22, 2024.***Letter states reason for recall, health risk and action to take:*** Please review this letter with your Medical Director to determine the appropriate course of***action, including for any previously generated results, if applicable.*** If you are a distributor, please ensure your customers receive this UMDC letter.*** To avoid the issue, please follow one of the options below:***o Option 1: After completing a patient test, either log out or return to the home***page before proceeding to the next patient test by pressing Close and Transmit ***button.***o Option 2: Before starting a new test, follow the steps outlined in Section 6.19*** Patient ID Lookup of the epoc System Manual with NXS Host and confirm the***correct patient demographics. It is important to verify the demographics, as some***of this data may be used for calculated analytes. If the demographics are***incorrect, delete and re-enter the Patient ID and perform the PPID Lookup again.*** Complete and return the Field Correction Effectiveness Check Form attached to this letter***within 30 days.*** Adverse reactions or quality problems experienced with the use of the affected products listed***in the above table may be reported to the FDA s MedWatch Adverse Event Reporting program***either online, by regular mail or by fax.*** Please retain this letter with your epoc System Manual with NXS Host or place the letter near***the affected instrument for reference. Please also forward this letter to those who may have***received this product.***For any questions, you may call Siemens Healthineers Technical Solutions Center at***877-229-3711, available daily from 7:00 AM to 11:00 PM EST.******The issue will be resolved in the | SoftwareDesign | • Software malfunction assigns incorrect patient name to blood measurement if user fails to verify, leading to incorrect actions. • Incorrect patient information saved with test result can lead to misdiagnosis and/or mistreatment, resulting in incorrect actions. • Failure to verify patient name can cause system to produce inaccurate results, leading to incorrect actions. | • Please review this letter with your Medical Director to determine the appropriate course of actions. • If you are a distributor, please ensure your customers receive this UMDC letter and perform the necessary actions. • To avoid the issue, please follow one of the options below and take the necessary actions: • Take actions by logging out or returning to the home page before proceeding to the next patient test by pressing the Close and Transmit button (Option 1). • Take actions by following the steps outlined in Section 6.19 Patient ID Lookup of the epoc System Manual with NXS Host and confirming the correct patient demographics (Option 2). • Take actions by completing and returning the Field Correction Effectiveness Check Form attached to this letter within the required timeframe. • Take actions by reporting any adverse reactions or quality problems experienced with the use of the affected products to the FDA's MedWatch Adverse Event Reporting program. • Take actions by retaining this letter with your epoc System Manual with NXS Host or placing the letter near the affected instrument for reference.
epoc Blood Analysis System | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics is recalling their epoc Blood Analysis System because of a software malfunction that occurs when specific criteria are met that can result in the incorrect patient name being assigned to a blood measurement by the device if the user of the device does not ensure verification of the correct patient name. The software malfunction has the potential to impact patient safety via misdiagnosis and/or mistreatment by a clinician if the wrong patient information is saved with the test result from the subject device. | Action Siemens issued Urgent Medical Device Correction (UMDC) letter to U.S. customers by FedEx beginning on November 22, 2024. Urgent Field Safety Notice (UFSN) was issued to OUS Siemens reporting sites with affected customers, for distribution and regional***reporting to their local competent authorities beginning on November 22, 2024.***Letter states reason for recall, health risk and action to take:*** Please review this letter with your Medical Director to determine the appropriate course of***action, including for any previously generated results, if applicable.*** If you are a distributor, please ensure your customers receive this UMDC letter.*** To avoid the issue, please follow one of the options below:***o Option 1: After completing a patient test, either log out or return to the home***page before proceeding to the next patient test by pressing Close and Transmit ***button.***o Option 2: Before starting a new test, follow the steps outlined in Section 6.19*** Patient ID Lookup of the epoc System Manual with NXS Host and confirm the***correct patient demographics. It is important to verify the demographics, as some***of this data may be used for calculated analytes. If the demographics are***incorrect, delete and re-enter the Patient ID and perform the PPID Lookup again.*** Complete and return the Field Correction Effectiveness Check Form attached to this letter***within 30 days.*** Adverse reactions or quality problems experienced with the use of the affected products listed***in the above table may be reported to the FDA s MedWatch Adverse Event Reporting program***either online, by regular mail or by fax.*** Please retain this letter with your epoc System Manual with NXS Host or place the letter near***the affected instrument for reference. Please also forward this letter to those who may have***received this product.***For any questions, you may call Siemens Healthineers Technical Solutions Center at***877-229-3711, available daily from 7:00 AM to 11:00 PM EST.******The issue will be resolved in the | SoftwareDesign | • Incorrect patient name assignment due to software malfunction, leading to potential misdiagnosis and/or mistreatment actions. • Failure of the device to ensure correct patient name verification, resulting in incorrect patient information being saved with test results actions. | • Review this letter with your Medical Director to determine the appropriate course of actions. • Follow one of the options to avoid the issue: Option 1, log out or return to the home page before proceeding to the next patient test, or Option 2, follow the patient ID lookup steps outlined in the epoc System Manual and confirm correct patient demographics, taking actions to verify and correct demographics if necessary. • Complete and return the Field Correction Effectiveness Check Form within the given timeframe, taking necessary actions to ensure timely submission. • Report any adverse reactions or quality problems experienced with the use of the affected products to the FDA’s MedWatch Adverse Event Reporting program, taking actions to report incidents promptly. • Retain this letter with your epoc System Manual or place it near the affected instrument for reference, taking actions to maintain accessible records. • Forward this letter to those who may have received this product, taking actions to ensure awareness. • Contact Siemens Healthineers Technical Solutions Center for any questions, taking actions to seek support when needed.
epoc Blood Analysis System | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics is recalling their epoc Blood Analysis System because of a software malfunction that occurs when specific criteria are met that can result in the incorrect patient name being assigned to a blood measurement by the device if the user of the device does not ensure verification of the correct patient name. The software malfunction has the potential to impact patient safety via misdiagnosis and/or mistreatment by a clinician if the wrong patient information is saved with the test result from the subject device. | Action Siemens issued Urgent Medical Device Correction (UMDC) letter to U.S. customers by FedEx beginning on November 22, 2024. Urgent Field Safety Notice (UFSN) was issued to OUS Siemens reporting sites with affected customers, for distribution and regional***reporting to their local competent authorities beginning on November 22, 2024.***Letter states reason for recall, health risk and action to take:*** Please review this letter with your Medical Director to determine the appropriate course of***action, including for any previously generated results, if applicable.*** If you are a distributor, please ensure your customers receive this UMDC letter.*** To avoid the issue, please follow one of the options below:***o Option 1: After completing a patient test, either log out or return to the home***page before proceeding to the next patient test by pressing Close and Transmit ***button.***o Option 2: Before starting a new test, follow the steps outlined in Section 6.19*** Patient ID Lookup of the epoc System Manual with NXS Host and confirm the***correct patient demographics. It is important to verify the demographics, as some***of this data may be used for calculated analytes. If the demographics are***incorrect, delete and re-enter the Patient ID and perform the PPID Lookup again.*** Complete and return the Field Correction Effectiveness Check Form attached to this letter***within 30 days.*** Adverse reactions or quality problems experienced with the use of the affected products listed***in the above table may be reported to the FDA s MedWatch Adverse Event Reporting program***either online, by regular mail or by fax.*** Please retain this letter with your epoc System Manual with NXS Host or place the letter near***the affected instrument for reference. Please also forward this letter to those who may have***received this product.***For any questions, you may call Siemens Healthineers Technical Solutions Center at***877-229-3711, available daily from 7:00 AM to 11:00 PM EST.******The issue will be resolved in the | SoftwareDesign | • Software malfunction assigns incorrect patient name to blood measurement if user does not verify, leading to potential misdiagnosis and mistreatment actions. • Incorrect patient information is saved with test result, compromising patient safety and requiring corrective actions. • Device failure to ensure correct patient name assignment results in potential harm to patients and necessitates immediate actions. | • Review this letter with your Medical Director to determine the appropriate course of actions. • Take one of the following actions to avoid the issue: • Perform actions as follows: log out or return to the home page before proceeding to the next patient test by pressing Close and Transmit button, or follow the steps outlined in Section 6.19 Patient ID Lookup of the epoc System Manual with NXS Host and confirm the correct patient demographics. • Verify the demographics, and if incorrect, delete and re-enter the Patient ID and perform the PPID Lookup again as part of your actions. • Complete and return the Field Correction Effectiveness Check Form attached to this letter as part of your actions. • Report any adverse reactions or quality problems experienced with the use of the affected products to the FDA’s MedWatch Adverse Event Reporting program as part of your actions.
epoc Blood Analysis System | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics is recalling their epoc Blood Analysis System because of a software malfunction that occurs when specific criteria are met that can result in the incorrect patient name being assigned to a blood measurement by the device if the user of the device does not ensure verification of the correct patient name. The software malfunction has the potential to impact patient safety via misdiagnosis and/or mistreatment by a clinician if the wrong patient information is saved with the test result from the subject device. | Action Siemens issued Urgent Medical Device Correction (UMDC) letter to U.S. customers by FedEx beginning on November 22, 2024. Urgent Field Safety Notice (UFSN) was issued to OUS Siemens reporting sites with affected customers, for distribution and regional***reporting to their local competent authorities beginning on November 22, 2024.***Letter states reason for recall, health risk and action to take:*** Please review this letter with your Medical Director to determine the appropriate course of***action, including for any previously generated results, if applicable.*** If you are a distributor, please ensure your customers receive this UMDC letter.*** To avoid the issue, please follow one of the options below:***o Option 1: After completing a patient test, either log out or return to the home***page before proceeding to the next patient test by pressing Close and Transmit ***button.***o Option 2: Before starting a new test, follow the steps outlined in Section 6.19*** Patient ID Lookup of the epoc System Manual with NXS Host and confirm the***correct patient demographics. It is important to verify the demographics, as some***of this data may be used for calculated analytes. If the demographics are***incorrect, delete and re-enter the Patient ID and perform the PPID Lookup again.*** Complete and return the Field Correction Effectiveness Check Form attached to this letter***within 30 days.*** Adverse reactions or quality problems experienced with the use of the affected products listed***in the above table may be reported to the FDA s MedWatch Adverse Event Reporting program***either online, by regular mail or by fax.*** Please retain this letter with your epoc System Manual with NXS Host or place the letter near***the affected instrument for reference. Please also forward this letter to those who may have***received this product.***For any questions, you may call Siemens Healthineers Technical Solutions Center at***877-229-3711, available daily from 7:00 AM to 11:00 PM EST.******The issue will be resolved in the | SoftwareDesign | • Incorrect patient name assignment actions due to software malfunction when specific criteria are met. • Potential for misdiagnosis actions due to incorrect patient information saved with test results. • Risk of mistreatment actions by clinicians relying on incorrect patient data. | • Please review this letter with your Medical Director to determine the appropriate course of actions. • If you are a distributor, please ensure your customers receive this UMDC letter and take the necessary actions. • To avoid the issue, please follow one of the options below: • Option 1: Take the actions of logging out or returning to the home page before proceeding to the next patient test by pressing Close and Transmit button. • Option 2: Take the actions outlined in Section 6.19 Patient ID Lookup of the epoc System Manual with NXS Host and confirm the correct patient demographics. • Complete and return the Field Correction Effectiveness Check Form attached to this letter within the specified timeframe and take the necessary actions. • For any questions, take the actions of calling Siemens Healthineers Technical Solutions Center.
epoc Blood Analysis System | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics is recalling their epoc Blood Analysis System because of a software malfunction that occurs when specific criteria are met that can result in the incorrect patient name being assigned to a blood measurement by the device if the user of the device does not ensure verification of the correct patient name. The software malfunction has the potential to impact patient safety via misdiagnosis and/or mistreatment by a clinician if the wrong patient information is saved with the test result from the subject device. | Action Siemens issued Urgent Medical Device Correction (UMDC) letter to U.S. customers by FedEx beginning on November 22, 2024. Urgent Field Safety Notice (UFSN) was issued to OUS Siemens reporting sites with affected customers, for distribution and regional***reporting to their local competent authorities beginning on November 22, 2024.***Letter states reason for recall, health risk and action to take:*** Please review this letter with your Medical Director to determine the appropriate course of***action, including for any previously generated results, if applicable.*** If you are a distributor, please ensure your customers receive this UMDC letter.*** To avoid the issue, please follow one of the options below:***o Option 1: After completing a patient test, either log out or return to the home***page before proceeding to the next patient test by pressing Close and Transmit ***button.***o Option 2: Before starting a new test, follow the steps outlined in Section 6.19*** Patient ID Lookup of the epoc System Manual with NXS Host and confirm the***correct patient demographics. It is important to verify the demographics, as some***of this data may be used for calculated analytes. If the demographics are***incorrect, delete and re-enter the Patient ID and perform the PPID Lookup again.*** Complete and return the Field Correction Effectiveness Check Form attached to this letter***within 30 days.*** Adverse reactions or quality problems experienced with the use of the affected products listed***in the above table may be reported to the FDA s MedWatch Adverse Event Reporting program***either online, by regular mail or by fax.*** Please retain this letter with your epoc System Manual with NXS Host or place the letter near***the affected instrument for reference. Please also forward this letter to those who may have***received this product.***For any questions, you may call Siemens Healthineers Technical Solutions Center at***877-229-3711, available daily from 7:00 AM to 11:00 PM EST.******The issue will be resolved in the | SoftwareDesign | • Software malfunction that assigns incorrect patient name to blood measurement if user does not verify correct patient name, leading to actions that can result in misdiagnosis and/or mistreatment. • Incorrect patient information being saved with test result, leading to actions that compromise patient safety. • Potential for incorrect patient data to influence clinician's actions, leading to actions that can cause harm to patients. | • Please review this letter with your Medical Director to determine the appropriate course of actions. • If you are a distributor, please ensure your customers receive this letter and take necessary actions. • To avoid the issue, take one of the following actions: • Follow Option 1: log out or return to the home page before proceeding to the next patient test by pressing Close and Transmit button. • Follow Option 2: complete the steps outlined in Section 6.19 Patient ID Lookup of the epoc System Manual with NXS Host and confirm the correct patient demographics. • Complete and return the Field Correction Effectiveness Check Form attached to this letter. • Verify the demographics, as some of this data may be used for calculated analytes, and take corrective actions. • If the demographics are incorrect, delete and re-enter the Patient ID and perform the PPID Lookup again. • Report any adverse reactions or quality problems experienced with the use of the affected products to the FDA's MedWatch Adverse Event Reporting program. • Retain this letter with your epoc System Manual with NXS Host or place the letter near the affected instrument for reference. • Forward this letter to those who may have received this product. • For any questions, call Siemens Healthineers Technical Solutions Center at 877-229-3711 and take necessary actions.
epoc Blood Analysis System | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics is recalling their epoc Blood Analysis System because of a software malfunction that occurs when specific criteria are met that can result in the incorrect patient name being assigned to a blood measurement by the device if the user of the device does not ensure verification of the correct patient name. The software malfunction has the potential to impact patient safety via misdiagnosis and/or mistreatment by a clinician if the wrong patient information is saved with the test result from the subject device. | Action Siemens issued Urgent Medical Device Correction (UMDC) letter to U.S. customers by FedEx beginning on November 22, 2024. Urgent Field Safety Notice (UFSN) was issued to OUS Siemens reporting sites with affected customers, for distribution and regional***reporting to their local competent authorities beginning on November 22, 2024.***Letter states reason for recall, health risk and action to take:*** Please review this letter with your Medical Director to determine the appropriate course of***action, including for any previously generated results, if applicable.*** If you are a distributor, please ensure your customers receive this UMDC letter.*** To avoid the issue, please follow one of the options below:***o Option 1: After completing a patient test, either log out or return to the home***page before proceeding to the next patient test by pressing Close and Transmit ***button.***o Option 2: Before starting a new test, follow the steps outlined in Section 6.19*** Patient ID Lookup of the epoc System Manual with NXS Host and confirm the***correct patient demographics. It is important to verify the demographics, as some***of this data may be used for calculated analytes. If the demographics are***incorrect, delete and re-enter the Patient ID and perform the PPID Lookup again.*** Complete and return the Field Correction Effectiveness Check Form attached to this letter***within 30 days.*** Adverse reactions or quality problems experienced with the use of the affected products listed***in the above table may be reported to the FDA s MedWatch Adverse Event Reporting program***either online, by regular mail or by fax.*** Please retain this letter with your epoc System Manual with NXS Host or place the letter near***the affected instrument for reference. Please also forward this letter to those who may have***received this product.***For any questions, you may call Siemens Healthineers Technical Solutions Center at***877-229-3711, available daily from 7:00 AM to 11:00 PM EST.******The issue will be resolved in the | SoftwareDesign | • Software malfunction can result in incorrect patient name being assigned to a blood measurement if user does not verify correct patient name, leading to actions that compromise patient safety. • Incorrect patient information being saved with test result can lead to misdiagnosis and/or mistreatment by a clinician, requiring corrective actions. | • Please review this letter with your Medical Director to determine the appropriate course of actions. • If you are a distributor, please ensure your customers receive this letter and take the necessary actions. • To avoid the issue, follow one of the options below and take the specified actions: • Option 1: Log out or return to the home page before proceeding to the next patient test by pressing Close and Transmit button, taking the required actions. • Option 2: Follow the steps outlined in Section 6.19 Patient ID Lookup of the epoc System Manual with NXS Host and confirm the correct patient demographics, taking the necessary actions. • Complete and return the Field Correction Effectiveness Check Form within the specified timeframe, taking the necessary actions. • Report any adverse reactions or quality problems experienced with the use of the affected products to the FDA's MedWatch Adverse Event Reporting program, taking the necessary actions. • Retain this letter with your epoc System Manual with NXS Host or place the letter near the affected instrument for reference, taking the necessary actions. • Forward this letter to those who may have received this product, taking the necessary actions. • For any questions, call Siemens Healthineers Technical Solutions Center and take the necessary actions.
epoc Blood Analysis System | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics is recalling their epoc Blood Analysis System because of a software malfunction that occurs when specific criteria are met that can result in the incorrect patient name being assigned to a blood measurement by the device if the user of the device does not ensure verification of the correct patient name. The software malfunction has the potential to impact patient safety via misdiagnosis and/or mistreatment by a clinician if the wrong patient information is saved with the test result from the subject device. | Action Siemens issued Urgent Medical Device Correction (UMDC) letter to U.S. customers by FedEx beginning on November 22, 2024. Urgent Field Safety Notice (UFSN) was issued to OUS Siemens reporting sites with affected customers, for distribution and regional***reporting to their local competent authorities beginning on November 22, 2024.***Letter states reason for recall, health risk and action to take:*** Please review this letter with your Medical Director to determine the appropriate course of***action, including for any previously generated results, if applicable.*** If you are a distributor, please ensure your customers receive this UMDC letter.*** To avoid the issue, please follow one of the options below:***o Option 1: After completing a patient test, either log out or return to the home***page before proceeding to the next patient test by pressing Close and Transmit ***button.***o Option 2: Before starting a new test, follow the steps outlined in Section 6.19*** Patient ID Lookup of the epoc System Manual with NXS Host and confirm the***correct patient demographics. It is important to verify the demographics, as some***of this data may be used for calculated analytes. If the demographics are***incorrect, delete and re-enter the Patient ID and perform the PPID Lookup again.*** Complete and return the Field Correction Effectiveness Check Form attached to this letter***within 30 days.*** Adverse reactions or quality problems experienced with the use of the affected products listed***in the above table may be reported to the FDA s MedWatch Adverse Event Reporting program***either online, by regular mail or by fax.*** Please retain this letter with your epoc System Manual with NXS Host or place the letter near***the affected instrument for reference. Please also forward this letter to those who may have***received this product.***For any questions, you may call Siemens Healthineers Technical Solutions Center at***877-229-3711, available daily from 7:00 AM to 11:00 PM EST.******The issue will be resolved in the | SoftwareDesign | • Software malfunction assigns incorrect patient name to blood measurement if user does not verify, leading to potential misdiagnosis and mistreatment actions. • Incorrect patient information is saved with test results, impacting patient safety through misdiagnosis and mistreatment actions. | • Please review this letter with your Medical Director to determine the appropriate course of actions, including for any previously generated results, if applicable. • To avoid the issue, take one of the following actions: • • Option 1: Log out or return to the home page before proceeding to the next patient test by pressing Close and Transmit button. • • Option 2: Follow the steps outlined in Section 6.19 Patient ID Lookup of the epoc System Manual with NXS Host and confirm the correct patient demographics. • Complete and return the Field Correction Effectiveness Check Form within 30 days, which includes reporting any adverse reactions or quality problems experienced with the use of the affected products. • For any questions, call Siemens Healthineers Technical Solutions Center, available daily from 7:00 AM to 11:00 PM EST.
Patient Information Center iX | Philips North America | During extended operation of the MX40 with the PIC iX, the DHCP (Dynamic Host Configuration Protocol) lease can expire on the MX40. If the DHCP lease expires while the device is offline, upon reconnecting with the PIC iX, a new IP address will be assigned to the MX40. The PIC iX will not recognize the new IP address after reconnecting and therefore the expecting settings sync fails. | Action An IMPORTANT PRODUCT NOTICE dated 4/19/23 was sent to customers.******The actions that you as a customer can take to minimize the effect of the problem***This communication should be shared with all clinical staff to review and understand. Place this Important Product Notice with the documentation of the PIC iX Instructions for Use.******Additional Actions Dependent on Normal Workflow:***1. Workf/ow A: MX40 Routinely Cleared from Sector:***If the MX40 is routinely cleared from the sector upon discharge of a patient then no additional action is required.***2. Workflow B: MX40 not Routinely Cleared from Sector:***Option 1: Change Default Settings to Clear Sector Automatically***The PIC iX can be setup to routinely clear the sector by executing the following:***a. This can be set as the default as needed by selecting the "Clear Sector" selection and***the "Clear unlocked Telemetry Devices from bed" in the "Configuration, Patient Management/Workflow", as highlighted in the image below:******Option 2: Clear MX40 from Sector Manually***The MX40 can be cleared from the sector manually by following the steps below:***a. Identify the sector to be cleared from the PIC iX. These steps may be repeated for all***sectors in use if they are not routinely cleared.***b. Select the "Manage Patient" button and the "Manage Patient" window will display.***c. Select the "Clear Sector'' button and the discharge window will display.***d. Check the box next to the "Clear Sector'' option, as shown in the example Image 4***below:******e. Selecting the "Cancel" button does not change the sector assignment. Finalize the***process by selecting the " Discharge" button.***3. Workflow C: Clear Sector Option Not Available:***In some configurations, the clear sector function will not be available. If this occurs, please contact the local response center for guidance on how to proceed.***4. The actions planned by Philips to correct the problem***A Philips representative will reach out to you to arrange for a software upgrade to your device. This upgrade wi | SoftwareDesign | • DHCP lease expiration on the MX40 causes IP address changes, leading to unexpected actions. • The PIC iX fails to recognize the new IP address, resulting in failed settings sync actions. • Device offline and reconnect issues trigger IP address changes, causing unexpected actions. | • To minimize the effect of the problem, actions you can take include reviewing and understanding this communication and sharing it with all clinical staff. • Actions to take dependent on normal workflow are: • Actions for Workflow A: no additional action is required if the MX40 is routinely cleared from the sector upon discharge of a patient. • Actions for Workflow B: Option 1, change default settings to clear sector automatically by selecting "Clear Sector" and "Clear unlocked Telemetry Devices from bed" in the "Configuration, Patient Management/Workflow". • Actions for Workflow B: Option 2, clear MX40 from sector manually by identifying the sector, selecting "Manage Patient" button, selecting "Clear Sector" button, checking the "Clear Sector" option, and finalizing the process by selecting "Discharge" button. • Actions for Workflow C: contact the local response center for guidance if the clear sector function is not available.
Beckman Coulter DxI 9000 Access Immunoassay Analyzer | Beckman Coulter, Inc. | When the DxI 9000 Access Immunoassay Analyzer with system software version 1.16.2 and prior is configured to automatically request a calibration order, the instrument can process the order with expired calibrator material. If the calibration curve is generated with an expired calibration material lot, this may lead to a hazardous situation of erroneously high or low patient results reported to the physicians. | Action The recalling firm issued letters dated 4/12/2024 on 4/18/2024 via first class mail. The letter described the issue, the impact, and actions to be taken, which included:******(a) Verify the calibrator material expiration date before ordering a calibration on the analyzer;******(b) Cancel calibration orders after the calibrator lot material expires. Instructions for canceling a calibration order are provided in the device help system or in the Calibration section of the IFU;******(c) Repeat any test results generated from a calibration curve that used expired calibrator lot material.******The resolution provided was that the firm is investigating the root cause of the issue and will implement the correction with a future software release. A service representative will contact the consignee to schedule the software upgrade when it is available.******The letter is to be shared with laboratory staff and retained as part of the laboratory Quality System documentation. If the device has been forwarded to another laboratory, the consignee is to provide a copy of the letter to that laboratory.******A Response Form was enclosed for completion and return within 10 days to assure the information has been received. | SoftwareDesign | • The instrument can process calibration orders with expired calibrator material, leading to actions that generate inaccurate results. • Calibration curves generated with expired material can produce erroneously high or low patient results, leading to actions that compromise patient safety. | • To address the issue, take the following actions: verify the calibrator material expiration date before ordering a calibration on the analyzer; cancel calibration orders after the calibrator lot material expires; repeat any test results generated from a calibration curve that used expired calibrator lot material. • Take actions to cancel a calibration order by following the instructions provided in the device help system or in the Calibration section of the IFU. • Share the letter with laboratory staff and take actions to retain it as part of the laboratory Quality System documentation. • Take actions to provide a copy of the letter to another laboratory if the device has been forwarded to them.
Cardiosave Hybrid IABP | Datascope Corp. | Firm has developed a software correction (Revision D.00) outside the US to mitigate product problems that pose a risk of hemodynamic instability and failure of transmission of clinical data in Cardiosave Hybrid and Cardiosave Rescue Intra-Aortic Balloon Pump devices. | Action Consignees were mailed an URGENT FIELD SAFETY NOTICE dated 8/5/2024. This notification informs consignees of the deployment of previously installed software to mitigate previously identified issues in Cardiosave Hybrid and/or Cardiosave Rescue Intra-Aortic Balloon Pump devices. Consignees are asked to examine their inventory immediately to identify if they have any Cardiosave Hybrid and/or Rescue IABP devices, and to identify what software version is currently installed in devices. Devices with software revision D.00 or D.01 have been corrected. If consignee units have a software version lower than D.00 or D.01 (e.g.: B.XX or C.XX) the device still requires the software update and has not yet been corrected. If consignee units do not have the corrected software revision installed, they are to continue to follow the instructions as listed int he original recall notification until the affected device(s) has been updated. Consignees are to complete and return the provided Response Form to the provided email address. Datascope/Getinge service representatives will contact consignees to schedule installation of software, if needed. Consignees are to forward the recall notice to all device users in their facility and to any customers to whom the device was shipped. | SoftwareDesign | • Risk of hemodynamic instability actions through Intra-Aortic Balloon Pump devices. • Failure of transmission of clinical data actions in Cardiosave Hybrid and Cardiosave Rescue devices. | • Examine your inventory immediately to identify if you have any Cardiosave Hybrid and/or Rescue IABP devices, and identify what software version is currently installed in devices, as these actions are required. • If your devices have a software version lower than D.00 or D.01, take actions to continue following the instructions as listed in the original recall notification until the affected device(s) has been updated. • Complete and return the provided Response Form to the provided email address, as these actions are necessary. • Forward the recall notice to all device users in your facility and to any customers to whom the device was shipped, and take actions to ensure they are informed. • Wait for Datascope/Getinge service representatives to contact you to schedule installation of software, if needed, and take actions to facilitate the installation process.
Cardiosave Hybrid IABP | Datascope Corp. | Firm has developed a software correction (Revision D.01) for outside the US to mitigate product problems that pose a risk of hemodynamic instability and failure of transmission of clinical data in Cardiosave Hybrid and Cardiosave Rescue Intra-Aortic Balloon Pump devices. | Action Consignees were mailed an URGENT FIELD SAFETY NOTICE dated 8/5/2024. This notification informs consignees of the deployment of previously installed software to mitigate previously identified issues in Cardiosave Hybrid and/or Cardiosave Rescue Intra-Aortic Balloon Pump devices. Consignees are asked to examine their inventory immediately to identify if they have any Cardiosave Hybrid and/or Rescue IABP devices, and to identify what software version is currently installed in devices. Devices with software revision D.00 or D.01 have been corrected. If consignee units have a software version lower than D.00 or D.01 (e.g.: B.XX or C.XX) the device still requires the software update and has not yet been corrected. If consignee units do not have the corrected software revision installed, they are to continue | SoftwareDesign | • Risk of hemodynamic instability actions. • Failure of transmission of clinical data actions. | • Examine your inventory immediately to identify if you have any Cardiosave Hybrid and/or Cardiosave Rescue IABP devices and identify what software version is currently installed in devices as part of the required actions. • Take actions to check if your device has software revision D.00 or D.01, as these have been corrected. • If your device has a software version lower than D.00 or D.01 (e.g.: B.XX or C.XX), take actions to update the software as it is still required. • Continue to take actions if your device does not have the corrected software revision installed.
Cardiosave Rescue IABP | Datascope Corp. | Firm has developed a software correction (Revision D.00) outside the US to mitigate product problems that pose a risk of hemodynamic instability and failure of transmission of clinical data in Cardiosave Hybrid and Cardiosave Rescue Intra-Aortic Balloon Pump devices. | Action Consignees were mailed an URGENT FIELD SAFETY NOTICE dated 8/5/2024. This notification informs consignees of the deployment of previously installed software to mitigate previously identified issues in Cardiosave Hybrid and/or Cardiosave Rescue Intra-Aortic Balloon Pump devices. Consignees are asked to examine their inventory immediately to identify if they have any Cardiosave Hybrid and/or Rescue IABP devices, and to identify what software version is currently installed in devices. Devices with software revision D.00 or D.01 have been corrected. If consignee units have a software version lower than D.00 or D.01 (e.g.: B.XX or C.XX) the device still requires the software update and has not yet been corrected. If consignee units do not have the corrected software revision installed, they are to continue to follow the instructions as listed int he original recall notification until the affected device(s) has been updated. Consignees are to complete and return the provided Response Form to the provided email address. Datascope/Getinge service representatives will contact consignees to schedule installation of software, if needed. Consignees are to forward the recall notice to all device users in their facility and to any customers to whom the device was shipped. | SoftwareDesign | • Risk of hemodynamic instability due to faulty device actions. • Failure of transmission of clinical data leading to inaccurate medical actions. | • Examine your inventory immediately to identify if you have any Cardiosave Hybrid and/or Rescue IABP devices, and to identify what software version is currently installed in devices as part of these actions. • Identify if your devices have a software version lower than D.00 or D.01 and take necessary actions. • Continue to follow the instructions as listed in the original recall notification until the affected device(s) has been updated if your devices do not have the corrected software revision installed, as part of these actions. • Complete and return the provided Response Form to the provided email address as part of these actions. • Forward the recall notice to all device users in your facility and to any customers to whom the device was shipped, and take necessary actions. • Await contact from Datascope/Getinge service representatives to schedule installation of software, if needed, as part of these actions.
Cardiosave Rescue IABP | Datascope Corp. | Firm has developed a software correction (Revision D.01) for outside the US to mitigate product problems that pose a risk of hemodynamic instability and failure of transmission of clinical data in Cardiosave Hybrid and Cardiosave Rescue Intra-Aortic Balloon Pump devices. | Action Consignees were mailed an URGENT FIELD SAFETY NOTICE dated 8/5/2024. This notification informs consignees of the deployment of previously installed software to mitigate previously identified issues in Cardiosave Hybrid and/or Cardiosave Rescue Intra-Aortic Balloon Pump devices. Consignees are asked to examine their inventory immediately to identify if they have any Cardiosave Hybrid and/or Rescue IABP devices, and to identify what software version is currently installed in devices. Devices with software revision D.00 or D.01 have been corrected. If consignee units have a software version lower than D.00 or D.01 (e.g.: B.XX or C.XX) the device still requires the software update and has not yet been corrected. If consignee units do not have the corrected software revision installed, they are to continue | SoftwareDesign | • Risk of hemodynamic instability actions • Failure of transmission of clinical data actions | • Examine your inventory immediately to identify if you have any Cardiosave Hybrid and/or Rescue IABP devices and what software version is currently installed in devices, as part of the necessary actions. • Identify if your devices have a software version lower than D.00 or D.01 (e.g.: B.XX or C.XX) and take actions to update the software as required. • Take actions to update the software if your device does not have the corrected software revision installed (D.00 or D.01). • Continue actions to mitigate previously identified issues in Cardiosave Hybrid and/or Rescue IABP devices until the software update is complete.
DxC 500 AU Clinical Chemistry Analyzer | Beckman Coulter Mishima K.K. | Due to a software issues, after the instrument processes 250 racks cumulatively, any subsequent racks with samples requiring rerun/reflex will be held in the Sample Handler's Buffer area and and error code "9000" will be reported. This issue may cause a delay of results. | Action during the week of 07/25/2024, the firm sent an "URGENT MEDICAL DEVICE RECALL" letter via email and postal mail to customers informing them that after the analyzer processes 250 racks cumulatively, any subsequent racks with samples requiring rerun/reflex will be held in the Sample Handler's Buffer area and an error code "9000" will be reported.******Customer are instructed:***. Monitor the rerun/reflex process if a rerun/reflex rack is held in the buffer area, manually retrieve the rack from the buffer area, manually order the rerun/reflex test and reload the sample onto the sample handler. Refer to Rerunning a Test section of the Instructions for Use on page 10-34.******Beckman Coulter has implemented the correction with software V1.4 which was released on 06/21/2024. A Beckman Coulter service representative will contact customers to schedule the software upgrade.******For any questions or assistance - contact the Beckman Coulter Customer Support via email https://www.beckmancoulter.com/en/support/contact-us | SoftwareDesign | • Software issues cause instrument to hold racks with rerun/reflex samples in buffer area after 250 racks are processed, requiring manual actions. • Error code "9000" is reported, indicating system malfunction, and necessitating corrective actions. • Delay of results is caused by instrument's inability to process samples smoothly, demanding intervention and recovery actions. | • To take the following actions: Monitor the rerun/reflex process if a rerun/reflex rack is held in the buffer area. • Take the actions to manually retrieve the rack from the buffer area. • Perform the actions to manually order the rerun/reflex test and reload the sample onto the sample handler.
DxI 9000 Access Immunoassay Analyzer | Beckman Coulter, Inc. | The reason for this recall is Beckman Coulter has determined that when a DxI 9000 Access Immunoassay Analyzer is connected to a Laboratory Automation System (LAS); is running with the software versions 1.17 and below; and a sample tube is sent from the LAS to DxI 9000 Access Immunoassay Analyzer, an out of sequence message will be sent from the Analyzer to the LAS. This issue may result in multiple error messages (i.e., a sample pick and place (PnP) error or a no result obtained (NRT) flag may be displayed) and subsequently a risk of delays in processing/reporting patient test results and patient treatment. | Action The recalling firm issued letters dated 5/16/2024 on 5/22/2024 via mail and email. The letter provided the issue, impact, and actions to take. The actions included: ******(1) If the issue occurs, initialize the system as prompted by the analyzer;***(2) Contact the Beckman Coulter representative and schedule a service visit if the laboratory has experienced sample PnP or aliquot pipettor error that is not resolved by initializing the system;***(3) If samples received an NRT flag, request the test again and reload the sample;***(4) Identify any sample tubes that were not processed and reload them on the Dxl 9000 Access Immunoassay Analyzer from the sample handler input area.******The letter informs the consignee that Beckman Coulter has identified the root cause of the issue and will implement the correction with a future software release. Beckman Coulter will contact the consignee to schedule the software upgrade when it is available. | SoftwareDesign | • The analyzer sends an "out of sequence" message to the Laboratory Automation System when a sample tube is sent from the LAS to the DxI 9000 Access Immunoassay Analyzer, resulting in error messages and delays in processing and reporting patient test results and treatment actions. • The malfunction leads to sample pick and place (PnP) error or a "no result obtained" (NRT) flag, causing delays in patient treatment actions. | • Take the following actions: If the issue occurs, initialize the system as prompted by the analyzer. • Take the following actions: Contact the Beckman Coulter representative and schedule a service visit if the laboratory has experienced sample PnP or aliquot pipettor error that is not resolved by initializing the system. • Take the following actions: If samples received an NRT flag, request the test again and reload the sample. • Take the following actions: Identify any sample tubes that were not processed and reload them on the Dxl 9000 Access Immunoassay Analyzer from the sample handler input area.
Philips | Philips North America Llc | Philips lntelliVue Patient Monitors MX400/450/500/550 Shipped or Upgraded with Incorrect Software Options.***Changes were made to reflect some options becoming standard capabilities for software-version N.x. In the process, the entitlements of software options MOS, M06, and M20 were removed for software versions L.x and M.x. As a result of these changes Patient Monitors manufactured with or updated to the latest versions of software L.x or M.x will not offer the capabilities> Mmonitors with software version K.x may experience this issue too, due to software version K being out of support and the devices being provided with L.x entitlements offered by software options MOS, M06, and M20. Potential for incorrect or delayed treatment for the patient, | Action Philips North America issued URGENT Medical Device Correction on 6/25/24 to End-Users. Philips Market Organizations will distribute the letters outside of the U.S.***Letter states reason for recall, health risk and action to take:******Use only default measurement capabilities, as indicated in Table 1 above, for monitoring purposes.***Pass this notice to all those who need to be aware within your organization or to any organization where affected devices have been potentially transferred.******A Philips representative will contact you to schedule a visit from a Philips Field Service Engineer who will reload the current device software to enable missing options (MOS, M06, M20).***If you need any further information, please contact your local Philips representative or call 1-800-722-9377. | SoftwareDesign | • Removal of software options MOS, M06, and M20 in software versions L.x and M.x may lead to incorrect or delayed treatment for the patient due to lack of necessary actions. • Inaccurate entitlements in software version K.x may cause devices to malfunction, requiring corrective actions. • Upgraded or shipped patient monitors with incorrect software options may not offer standard capabilities, necessitating remedial actions. | • Take actions by using only default measurement capabilities for monitoring purposes. • Pass this notice to all those who need to be aware within your organization or to any organization where affected devices have been potentially transferred. • Expect a Philips representative to contact you to schedule a visit from a Philips Field Service Engineer who will reload the current device software to enable missing options. • If you need further information, take actions by contacting your local Philips representative or call 1-800-722-9377.
Philips | Philips North America Llc | Philips lntelliVue Patient Monitors MX400/450/500/550 Shipped or Upgraded with Incorrect Software Options.***Changes were made to reflect some options becoming standard capabilities for software-version N.x. In the process, the entitlements of software options MOS, M06, and M20 were removed for software versions L.x and M.x. As a result of these changes Patient Monitors manufactured with or updated to the latest versions of software L.x or M.x will not offer the capabilities> Mmonitors with software version K.x may experience this issue too, due to software version K being out of support and the devices being provided with L.x entitlements offered by software options MOS, M06, and M20. Potential for incorrect or delayed treatment for the patient, | Action Philips North America issued URGENT Medical Device Correction on 6/25/24 to End-Users. Philips Market Organizations will distribute the letters outside of the U.S.***Letter states reason for recall, health risk and action to take:******Use only default measurement capabilities, as indicated in Table 1 above, for monitoring purposes.***Pass this notice to all those who need to be aware within your organization or to any organization where affected devices have been potentially transferred.******A Philips representative will contact you to schedule a visit from a Philips Field Service Engineer who will reload the current device software to enable missing options (MOS, M06, M20).***If you need any further information, please contact your local Philips representative or call 1-800-722-9377. | SoftwareDesign | • Incorrect software options led to removal of entitlements for software versions L.x and M.x, affecting patient monitors' capabilities, and requiring actions. • This issue may cause incorrect or delayed treatment for patients, necessitating immediate actions. • Software version K.x devices, no longer supported, may also experience this problem, demanding prompt actions. | • Take the following actions: use only default measurement capabilities for monitoring purposes, as indicated in Table 1 above. • Pass this notice to all those who need to be aware within your organization or to any organization where affected devices have been potentially transferred. • Expect a Philips representative to contact you to schedule a visit from a Philips Field Service Engineer who will reload the current device software to enable missing options. • If you need further information, take the actions to contact your local Philips representative or call 1-800-722-9377.
Philips | Philips North America Llc | Philips lntelliVue Patient Monitors MX400/450/500/550 Shipped or Upgraded with Incorrect Software Options.***Changes were made to reflect some options becoming standard capabilities for software-version N.x. In the process, the entitlements of software options MOS, M06, and M20 were removed for software versions L.x and M.x. As a result of these changes Patient Monitors manufactured with or updated to the latest versions of software L.x or M.x will not offer the capabilities> Mmonitors with software version K.x may experience this issue too, due to software version K being out of support and the devices being provided with L.x entitlements offered by software options MOS, M06, and M20. Potential for incorrect or delayed treatment for the patient, | Action Philips North America issued URGENT Medical Device Correction on 6/25/24 to End-Users. Philips Market Organizations will distribute the letters outside of the U.S.***Letter states reason for recall, health risk and action to take:******Use only default measurement capabilities, as indicated in Table 1 above, for monitoring purposes.***Pass this notice to all those who need to be aware within your organization or to any organization where affected devices have been potentially transferred.******A Philips representative will contact you to schedule a visit from a Philips Field Service Engineer who will reload the current device software to enable missing options (MOS, M06, M20).***If you need any further information, please contact your local Philips representative or call 1-800-722-9377. | SoftwareDesign | • Incorrect software options led to removal of entitlements for software versions L.x and M.x, affecting patient monitor capabilities and potential patient treatment actions. • Software versions L.x and M.x do not offer capabilities due to removed entitlements, potentially causing incorrect or delayed treatment actions. | • Take the following actions: Use only default measurement capabilities for monitoring purposes. • Pass this notice to all those who need to be aware within your organization or to any organization where affected devices have been potentially transferred. • Wait for a Philips representative to contact you to schedule a visit from a Philips Field Service Engineer who will reload the current device software to enable missing options. • If you need any further information, take the actions of contacting your local Philips representative or calling 1-800-722-9377.
Philips | Philips North America Llc | Philips lntelliVue Patient Monitors MX400/450/500/550 Shipped or Upgraded with Incorrect Software Options.***Changes were made to reflect some options becoming standard capabilities for software-version N.x. In the process, the entitlements of software options MOS, M06, and M20 were removed for software versions L.x and M.x. As a result of these changes Patient Monitors manufactured with or updated to the latest versions of software L.x or M.x will not offer the capabilities> Mmonitors with software version K.x may experience this issue too, due to software version K being out of support and the devices being provided with L.x entitlements offered by software options MOS, M06, and M20. Potential for incorrect or delayed treatment for the patient, | Action Philips North America issued URGENT Medical Device Correction on 6/25/24 to End-Users. Philips Market Organizations will distribute the letters outside of the U.S.***Letter states reason for recall, health risk and action to take:******Use only default measurement capabilities, as indicated in Table 1 above, for monitoring purposes.***Pass this notice to all those who need to be aware within your organization or to any organization where affected devices have been potentially transferred.******A Philips representative will contact you to schedule a visit from a Philips Field Service Engineer who will reload the current device software to enable missing options (MOS, M06, M20).***If you need any further information, please contact your local Philips representative or call 1-800-722-9377. | SoftwareDesign | • Incorrect software options led to removal of essential capabilities, affecting patient treatment actions. • Incompatible software versions (L.x, M.x, and K.x) caused inconsistent device performance, impacting patient safety actions. • Patient monitors shipped or upgraded with incorrect software options resulted in potential incorrect or delayed treatment actions. | • Use only default measurement capabilities for monitoring purposes, as indicated in Table 1, for necessary actions. • Pass this notice to all those who need to be aware within your organization or to any organization where affected devices have been potentially transferred for timely actions. • Schedule a visit from a Philips Field Service Engineer for actions to reload the current device software to enable missing options (MOS, M06, M20). • If you need any further information, contact your local Philips representative or call 1-800-722-9377 for necessary actions.
DxC 500 AU Clinical Chemistry Analyzer | Beckman Coulter Mishima K.K. | There are two software bug issues with the analyzer. Issue 1: When a dedicated rack is reserved for calibration, after a calibration test has been placed and completed, the same rack may be used to place both a patient order and a calibration order. When a patient sample and a non-barcoded calibrator tube is processed on the rack, the calibration result will be reported as a patient result. Issue 2: calibration with expired calibrator. They both may cause erroneous patient results, but the probability of serious adverse health consequences or death is unlikely. *** | Action During the week of June 20, 2024, the firm sent an "URGENT MEDICAL DEVICE RECALL" Letter via email/postal mail to customers informing them of two software related issues associated with use of the Clinical Chemistry Analyzer.******Issue 1:*** When a dedicated rack is reserved for calibration, after a calibration test has been placed and completed, the same rack may be used to place both a patient order and a calibration order. When a patient sample and a non-barcoded calibrator tube is processed on the rack, the calibration result will be reported as a patient result. This may cause a false low result or a false high result to be reported to a physician.******Issues 2:***When a calibration order is manually placed (calibrator lot is not expired) and then the calibrator expires before the calibrator material is placed onto the analyzer for processing, the analyzer will process the expired calibration lot and generate patient results that may be inaccurate, potentially causing a false low result or a false high result to be reported to a physician.******Customer are instructed to:***For Issue 1***-Follow the IFU Note in Assigning Racks or Calibrator IDs to Calibrators section on pg. 8-7.***-Alternatively, change calibrator Load Type to Bar Code Configuration and use barcoded containers for calibration.***-Refer to IFU Assigning Racks or Calibrator IDs to Calibrators section on pg. 8-6, and in the software user interface Calibration section, under Calibrations Due Now, review the Load Type configuration, rack ID and cup position set up for the calibrators.******Issue 2:***-Verify that the calibrator has not expired prior to loading the calibrator material onto the analyzer for processing.***-Cancel calibration orders after the calibrator lot material expires.***-Repeat any test results generated from a calibration curve that used expired calibrator lot material.******For questions/assistance-contact Customer Support Center at http://www.beckmancoulter.com | SoftwareDesign | • Analyzer allows a dedicated rack reserved for calibration to be reused for patient orders, leading to incorrect reporting of calibration results as patient results through improper actions. • Analyzer permits use of expired calibrators for calibration, resulting in potential erroneous patient results through faulty actions. | • For Issue 1, take actions by following the IFU Note in Assigning Racks or Calibrator IDs to Calibrators section, or change calibrator Load Type to Bar Code Configuration and use barcoded containers for calibration, and review the Load Type configuration, rack ID, and cup position set up for the calibrators. • For Issue 2, take actions by verifying that the calibrator has not expired prior to loading the calibrator material onto the analyzer for processing, cancel calibration orders after the calibrator lot material expires, and repeat any test results generated from a calibration curve that used expired calibrator lot material.
Siemens | Siemens Medical Solutions USA, Inc | If the Sensis documentation functionality is used during adding of once-per-study reporting events (e.g., Type 1 events, as further defined in the administrator manual), the possibility exists that the application could crash. | Action Siemens issued Urgent Medical Deice Correction letter (AX038/23/S-C) on 7/3/24 to Director of Diagnostic Imaging and/or Radiology Laboratory. Letter states reason for recall, health risk and action to take:***If the Sensis documentation functionality is used and the described scenario occurs, users should continue the exam by***opening a new study for the same patient without adding any Type 1 reporting events. You can compensate this by adding***free text comments (procedure notes) only and manually inserting the desired information in those comments.***Please ensure that patient treatment can be continued in other ways if there is any possible danger for the safety of the***patient.***How will the corrective action be implemented and what is the efficiency of the corrective actions?***Siemens will correct the error with a software update via Update Instruction AX010/23/S. The corrective action mitigates***the probability of occurrence of the issue. Based on our investigation you can continue to use your system/device.***Our service organization will contact you shortly to arrange a date to perform this corrective action. Please feel free to contact our service organization for an earlier appointment at 1-800-888-7436.******Please review, understand, and confirm acknowledgement of the information and instructions outlined in this Customer***Safety Advisory Notice. Please respond within 14 days of receiving this notification.***If you have received this notification by email via Adobe GigaSign, kindly sign with your digital signature to acknowledge that***you have read and understand the content provided.***If you have received this customer notification by delivery by the US Postal Service or by FedEx, please send an email to recallsandrefusals.team@siemens-healthineers.com acknowledging that you have read and understand the content provided | SoftwareDesign | • Potential application crash when using Sensis documentation functionality during adding of once-per-study reporting events, which may lead to loss of data or disruption of actions. • Possibility of system failure, affecting the operation and performance of actions. • Risk of system downtime, causing delay or hindrance in taking further actions. | • Take actions to continue the exam by opening a new study for the same patient without adding any Type 1 reporting events and compensate by adding free text comments (procedure notes) only and manually inserting the desired information in those comments to mitigate the health risk. • Take actions to ensure patient treatment can be continued in other ways if there is any possible danger for the safety of the patient. • Take actions to review, understand, and confirm acknowledgement of the information and instructions outlined in this Customer Safety Advisory Notice and respond by signing with a digital signature or sending an email to recallsandrefusals.team@siemens-healthineers.com within 14 days.
Siemens | Siemens Medical Solutions USA, Inc | If the Sensis documentation functionality is used during adding of once-per-study reporting events (e.g., Type 1 events, as further defined in the administrator manual), the possibility exists that the application could crash. | Action Siemens issued Urgent Medical Deice Correction letter (AX038/23/S-C) on 7/3/24 to Director of Diagnostic Imaging and/or Radiology Laboratory. Letter states reason for recall, health risk and action to take:***If the Sensis documentation functionality is used and the described scenario occurs, users should continue the exam by***opening a new study for the same patient without adding any Type 1 reporting events. You can compensate this by adding***free text comments (procedure notes) only and manually inserting the desired information in those comments.***Please ensure that patient treatment can be continued in other ways if there is any possible danger for the safety of the***patient.***How will the corrective action be implemented and what is the efficiency of the corrective actions?***Siemens will correct the error with a software update via Update Instruction AX010/23/S. The corrective action mitigates***the probability of occurrence of the issue. Based on our investigation you can continue to use your system/device.***Our service organization will contact you shortly to arrange a date to perform this corrective action. Please feel free to contact our service organization for an earlier appointment at 1-800-888-7436.******Please review, understand, and confirm acknowledgement of the information and instructions outlined in this Customer***Safety Advisory Notice. Please respond within 14 days of receiving this notification.***If you have received this notification by email via Adobe GigaSign, kindly sign with your digital signature to acknowledge that***you have read and understand the content provided.***If you have received this customer notification by delivery by the US Postal Service or by FedEx, please send an email to recallsandrefusals.team@siemens-healthineers.com acknowledging that you have read and understand the content provided | SoftwareDesign | • Crashing of the application during adding of once-per-study reporting events due to Sensis documentation functionality usage, affecting system operation and performance through actions. • Possibility of system crash when using Sensis documentation functionality with Type 1 events, impacting system safety and reliability through actions. | • If the Sensis documentation functionality is used and the described scenario occurs, take the following actions: open a new study for the same patient without adding any Type 1 reporting events, and compensate by adding free text comments (procedure notes) only and manually inserting the desired information in those comments. • Ensure patient treatment can be continued in other ways if there is any possible danger for the safety of the patient by taking the necessary actions. • Review, understand, and confirm acknowledgement of the information and instructions outlined in this Customer Safety Advisory Notice by taking the actions: respond within the given timeframe, and sign with your digital signature or send an email to acknowledge that you have read and understand the content provided. • Contact Siemens' service organization for an earlier appointment to arrange a date to perform the corrective action, and take actions to implement the software update via Update Instruction AX010/23/S.
Siemens | Siemens Medical Solutions USA, Inc | If the Sensis documentation functionality is used during adding of once-per-study reporting events (e.g., Type 1 events, as further defined in the administrator manual), the possibility exists that the application could crash. | Action Siemens issued Urgent Medical Deice Correction letter (AX038/23/S-C) on 7/3/24 to Director of Diagnostic Imaging and/or Radiology Laboratory. Letter states reason for recall, health risk and action to take:***If the Sensis documentation functionality is used and the described scenario occurs, users should continue the exam by***opening a new study for the same patient without adding any Type 1 reporting events. You can compensate this by adding***free text comments (procedure notes) only and manually inserting the desired information in those comments.***Please ensure that patient treatment can be continued in other ways if there is any possible danger for the safety of the***patient.***How will the corrective action be implemented and what is the efficiency of the corrective actions?***Siemens will correct the error with a software update via Update Instruction AX010/23/S. The corrective action mitigates***the probability of occurrence of the issue. Based on our investigation you can continue to use your system/device.***Our service organization will contact you shortly to arrange a date to perform this corrective action. Please feel free to contact our service organization for an earlier appointment at 1-800-888-7436.******Please review, understand, and confirm acknowledgement of the information and instructions outlined in this Customer***Safety Advisory Notice. Please respond within 14 days of receiving this notification.***If you have received this notification by email via Adobe GigaSign, kindly sign with your digital signature to acknowledge that***you have read and understand the content provided.***If you have received this customer notification by delivery by the US Postal Service or by FedEx, please send an email to recallsandrefusals.team@siemens-healthineers.com acknowledging that you have read and understand the content provided | SoftwareDesign | • Crash of the application occurs when Sensis documentation functionality is used during adding of once-per-study reporting events, which affects the system's operation and performance through faulty actions. • The crash poses a risk to data integrity and user experience due to faulty actions. | • Continue the exam by opening a new study for the same patient without adding any Type 1 reporting events and take actions by adding free text comments (procedure notes) only and manually inserting the desired information in those comments. • Ensure patient treatment can be continued in other ways if there is any possible danger for the safety of the patient and take actions accordingly. • Review, understand, and confirm acknowledgement of the information and instructions outlined in this Customer Safety Advisory Notice and take actions by responding within 14 days of receiving this notification.
Siemens | Siemens Medical Solutions USA, Inc | If the Sensis documentation functionality is used during adding of once-per-study reporting events (e.g., Type 1 events, as further defined in the administrator manual), the possibility exists that the application could crash. | Action Siemens issued Urgent Medical Deice Correction letter (AX038/23/S-C) on 7/3/24 to Director of Diagnostic Imaging and/or Radiology Laboratory. Letter states reason for recall, health risk and action to take:***If the Sensis documentation functionality is used and the described scenario occurs, users should continue the exam by***opening a new study for the same patient without adding any Type 1 reporting events. You can compensate this by adding***free text comments (procedure notes) only and manually inserting the desired information in those comments.***Please ensure that patient treatment can be continued in other ways if there is any possible danger for the safety of the***patient.***How will the corrective action be implemented and what is the efficiency of the corrective actions?***Siemens will correct the error with a software update via Update Instruction AX010/23/S. The corrective action mitigates***the probability of occurrence of the issue. Based on our investigation you can continue to use your system/device.***Our service organization will contact you shortly to arrange a date to perform this corrective action. Please feel free to contact our service organization for an earlier appointment at 1-800-888-7436.******Please review, understand, and confirm acknowledgement of the information and instructions outlined in this Customer***Safety Advisory Notice. Please respond within 14 days of receiving this notification.***If you have received this notification by email via Adobe GigaSign, kindly sign with your digital signature to acknowledge that***you have read and understand the content provided.***If you have received this customer notification by delivery by the US Postal Service or by FedEx, please send an email to recallsandrefusals.team@siemens-healthineers.com acknowledging that you have read and understand the content provided | SoftwareDesign | • Application crash actions may occur when using Sensis documentation functionality during adding of once-per-study reporting events, affecting system operation. | • Take actions: If the Sensis documentation functionality is used, continue the exam by opening a new study for the same patient without adding any Type 1 reporting events. • Take actions: Compensate by adding free text comments (procedure notes) only and manually inserting the desired information in those comments. • Take actions: Ensure patient treatment can be continued in other ways if there is any possible danger for the safety of the patient. • Take actions: Review, understand, and confirm acknowledgement of the information and instructions outlined in this Customer Safety Advisory Notice.
DxI 9000 Access Immunoassay Analyzer | Beckman Coulter, Inc. | Beckman Coulter is recalling their DxI 9000 Access Immunoassay Analyzers, an in vitro diagnostic device used for the quantitative, semi-quantitative, or qualitative determination of various analyte concentrations found in human body fluids, by correction. The reason for the recall is: Beckman Coulter has determined that the DxI 9000 Access Immunoassay analyzer accepts calibrator values scanned from multiple calibrator cards when an operator configures a new calibrator lot. The analyzer accepts the calibrator lot values from any scanned calibrator card, regardless of the assay test code or calibrator material lot number. The software is not validating the integrity of each calibrator barcode scanned to ensure they are all from the same calibrator card. This includes all software versions and all DxI 9000 instruments that are currently in the field. This issue does not impact calibrator material with 2D calibrator card barcodes. This may cause a failed calibration curve. The health risk to patients is dependent upon the analytes that could be impacted by the error, to include receiving delayed or erroneous results. ******The issue was identified by Beckman Coulter while conducting a field action on DxI 9000 Access Immunoassay Analyzers.*** *** | Action Beckman Coulter issued an URGENT MEDICAL DEVICE RECALL notice to its consignees on 04/18/2024 via USPS first class mail. The notice explained the issue, impact, and requested the following actions be taken: ***"x Confirm that the DxI 9000 Access Immunoassay Analyzer test name printed on the calibrator card matches the calibrator material name before you configure the calibrator lot.***x Scan each individual calibrator card and complete each step displayed on the DxI 9000 Access Immunoassay Analyzer user interface before you scan the next calibrator card.***x Review the expected values column on the calibration summary details page and confirm that values for that assay match the values printed on the calibration card. ***If a calibrator lot had incorrect calibrator levels configured take the following steps:***x Delete the impacted calibrator lot: Select Menu > Calibration Material List, select the desired calibration lot, and select Delete. ***x Reconfigure the correct calibrator lot as described in the Calibration section of the Instructions for Use (IFU).***x Rerun the calibration." ***"Your Beckman Coulter service representative will contact you to schedule the software upgrade when it is available. Please share this information with your laboratory staff and retain this notification as part of your laboratory Quality System documentation. If you have forwarded any of the affected product(s) listed above to another laboratory, please provide them a copy of this letter." | SoftwareDesign | • The analyzer accepts calibrator lot values from any scanned calibrator card, regardless of the assay test code or calibrator material lot number, which may cause failed calibration curves and impact calibration actions. • The software does not validate the integrity of each calibrator barcode scanned, leading to potential errors and affecting calibration actions. • This issue may cause patients to receive delayed or erroneous results, depending on the affected analytes, which can impact patient care actions. | • Confirm that the DxI 9000 Access Immunoassay Analyzer test name printed on the calibrator card matches the calibrator material name before you configure the calibrator lot actions. • Scan each individual calibrator card and complete each step displayed on the DxI 9000 Access Immunoassay Analyzer user interface before you scan the next calibrator card actions. • Review the expected values column on the calibration summary details page and confirm that values for that assay match the values printed on the calibration card actions. • If a calibrator lot had incorrect calibrator levels configured take the following actions: • Delete the impacted calibrator lot actions. • Reconfigure the correct calibrator lot as described in the Calibration section of the Instructions for Use (IFU) actions. • Rerun the calibration actions.
RayStation | RAYSEARCH LABORATORIES AB | Potential for reported SSD to be too high. | Action A Field Safety Notice for Medical Device Correction #133261, dated 3/28/24, was sent to consignees via email. Consignees were provided the instructions to use the Measure Tool to measure the distance in the patient view to make sure the reported SSD value is correct if it is to be used for patient setup or other critical tasks. If the SSD value is found to be incorrect, small changes to the isocenter, beam angles or the affected ROI's geometry can be made until the correct value is achieved. Consignees are to inform planning staff and all users of the workarounds provided in the Field Safety Notice. Consignees are to inspect their product and identify if installed units possess the affected software versions. Consignees are to reply to the recall notification confirming that they read and understood the Field Safety Notificaiton. RaySearch is to deploy a new version of RayStation/RayPlan to correct the issue, to be released in April 2024. The provided Field Safety Notice is to be maintained with affected units. Consignees with any questions are to contact quality@raysearchlabs.com. | SoftwareDesign | • Inaccurate SSD reporting actions may lead to incorrect system performance assessments. • Overreported SSD capacity actions can result in unintended system behavior or crashes. | • Follow the instructions provided in the Field Safety Notice to use the Measure Tool to validate the SSD value for patient setup and other critical tasks, and make necessary actions to correct any incorrect values. • Inform planning staff and all users of the workarounds provided in the Field Safety Notice and take necessary actions to ensure correct usage. • Inspect your product and identify if installed units possess the affected software versions, and take actions to correct any issues found. • Reply to the recall notification confirming that you read and understood the Field Safety Notice and take necessary actions as instructed. • Contact quality@raysearchlabs.com with any questions or concerns regarding the necessary actions.
RayStation | RAYSEARCH LABORATORIES AB | Potential for reported SSD to be too high. | Action A Field Safety Notice for Medical Device Correction #133261, dated 3/28/24, was sent to consignees via email. Consignees were provided the instructions to use the Measure Tool to measure the distance in the patient view to make sure the reported SSD value is correct if it is to be used for patient setup or other critical tasks. If the SSD value is found to be incorrect, small changes to the isocenter, beam angles or the affected ROI's geometry can be made until the correct value is achieved. Consignees are to inform planning staff and all users of the workarounds provided in the Field Safety Notice. Consignees are to inspect their product and identify if installed units possess the affected software versions. Consignees are to reply to the recall notification confirming that they read and understood the Field Safety Notificaiton. RaySearch is to deploy a new version of RayStation/RayPlan to correct the issue, to be released in April 2024. The provided Field Safety Notice is to be maintained with affected units. Consignees with any questions are to contact quality@raysearchlabs.com. | SoftwareDesign | • Inaccurate SSD reporting actions may lead to incorrect system assessment and decision-making. • Potential for reported SSD to be too high actions can result in false diagnosis or misinterpretation of system performance. | • Take actions to inspect your product and identify if installed units possess the affected software versions. • Take actions to use the Measure Tool to measure the distance in the patient view to ensure the reported SSD value is correct. • Take actions to inform planning staff and all users of the workarounds provided in the Field Safety Notice. • Take actions to reply to the recall notification confirming that you read and understood the Field Safety Notice. • Take actions to contact quality@raysearchlabs.com if you have any questions. • Take actions to maintain the provided Field Safety Notice with affected units.
RayStation | RAYSEARCH LABORATORIES AB | Potential for reported SSD to be too high. | Action A Field Safety Notice for Medical Device Correction #133261, dated 3/28/24, was sent to consignees via email. Consignees were provided the instructions to use the Measure Tool to measure the distance in the patient view to make sure the reported SSD value is correct if it is to be used for patient setup or other critical tasks. If the SSD value is found to be incorrect, small changes to the isocenter, beam angles or the affected ROI's geometry can be made until the correct value is achieved. Consignees are to inform planning staff and all users of the workarounds provided in the Field Safety Notice. Consignees are to inspect their product and identify if installed units possess the affected software versions. Consignees are to reply to the recall notification confirming that they read and understood the Field Safety Notificaiton. RaySearch is to deploy a new version of RayStation/RayPlan to correct the issue, to be released in April 2024. The provided Field Safety Notice is to be maintained with affected units. Consignees with any questions are to contact quality@raysearchlabs.com. | SoftwareDesign | • Inaccurate SSD reporting may lead to faulty actions. • Incorrectly reported SSD may cause system misconfiguration, affecting actions. • Potential for misleading system performance metrics that influence actions. | • Take actions to use the Measure Tool to measure the distance in the patient view to ensure the reported SSD value is correct if it is to be used for patient setup or other critical tasks. • Take actions to inform planning staff and all users of the workarounds provided in the Field Safety Notice. • Take actions to inspect your product and identify if installed units possess the affected software versions. • Take actions to reply to the recall notification confirming that you read and understood the Field Safety Notification. • Take actions to maintain the provided Field Safety Notice with affected units. • Take actions to contact quality@raysearchlabs.com if you have any questions.
RayStation | RAYSEARCH LABORATORIES AB | Potential for reported SSD to be too high. | Action A Field Safety Notice for Medical Device Correction #133261, dated 3/28/24, was sent to consignees via email. Consignees were provided the instructions to use the Measure Tool to measure the distance in the patient view to make sure the reported SSD value is correct if it is to be used for patient setup or other critical tasks. If the SSD value is found to be incorrect, small changes to the isocenter, beam angles or the affected ROI's geometry can be made until the correct value is achieved. Consignees are to inform planning staff and all users of the workarounds provided in the Field Safety Notice. Consignees are to inspect their product and identify if installed units possess the affected software versions. Consignees are to reply to the recall notification confirming that they read and understood the Field Safety Notificaiton. RaySearch is to deploy a new version of RayStation/RayPlan to correct the issue, to be released in April 2024. The provided Field Safety Notice is to be maintained with affected units. Consignees with any questions are to contact quality@raysearchlabs.com. | SoftwareDesign | • Inaccurate SSD reporting actions may lead to incorrect storage capacity assessments. • Potential overestimation of SSD capacity actions can result in data storage and retrieval issues. • Incorrect SSD reporting actions may impact system performance and reliability. | • Take actions to measure the distance in the patient view to ensure the reported SSD value is correct if it is to be used for patient setup or other critical tasks using the Measure Tool. • Take actions to make small changes to the isocenter, beam angles, or the affected ROI's geometry if the SSD value is found to be incorrect until the correct value is achieved. • Inform planning staff and all users of the workarounds provided in the Field Safety Notice as part of your actions. • Inspect your product and identify if installed units possess the affected software versions as part of your actions. • Reply to the recall notification confirming that you read and understood the Field Safety Notification as part of your actions. • Contact quality@raysearchlabs.com if you have any questions about the actions required.
RayStation | RAYSEARCH LABORATORIES AB | Potential for reported SSD to be too high. | Action A Field Safety Notice for Medical Device Correction #133261, dated 3/28/24, was sent to consignees via email. Consignees were provided the instructions to use the Measure Tool to measure the distance in the patient view to make sure the reported SSD value is correct if it is to be used for patient setup or other critical tasks. If the SSD value is found to be incorrect, small changes to the isocenter, beam angles or the affected ROI's geometry can be made until the correct value is achieved. Consignees are to inform planning staff and all users of the workarounds provided in the Field Safety Notice. Consignees are to inspect their product and identify if installed units possess the affected software versions. Consignees are to reply to the recall notification confirming that they read and understood the Field Safety Notificaiton. RaySearch is to deploy a new version of RayStation/RayPlan to correct the issue, to be released in April 2024. The provided Field Safety Notice is to be maintained with affected units. Consignees with any questions are to contact quality@raysearchlabs.com. | SoftwareDesign | • Incorrect SSD reporting actions may provide false information to users. • Inaccurate SSD reporting actions can lead to incorrect system performance assessments. • Potential system malfunction or failure actions may occur due to reliance on false SSD reports. | • Take actions to use the Measure Tool to measure the distance in the patient view to ensure the reported SSD value is correct for patient setup or other critical tasks. • Perform actions to make small changes to the isocenter, beam angles or the affected ROI's geometry if the SSD value is found to be incorrect. • Take actions to inform planning staff and all users of the workarounds provided in the Field Safety Notice. • Conduct actions to inspect your product and identify if installed units possess the affected software versions. • Take actions to reply to the recall notification confirming that you read and understood the Field Safety Notification. • Contact quality@raysearchlabs.com if you have any questions regarding the actions.
RayStation | RAYSEARCH LABORATORIES AB | Potential for reported SSD to be too high. | Action A Field Safety Notice for Medical Device Correction #133261, dated 3/28/24, was sent to consignees via email. Consignees were provided the instructions to use the Measure Tool to measure the distance in the patient view to make sure the reported SSD value is correct if it is to be used for patient setup or other critical tasks. If the SSD value is found to be incorrect, small changes to the isocenter, beam angles or the affected ROI's geometry can be made until the correct value is achieved. Consignees are to inform planning staff and all users of the workarounds provided in the Field Safety Notice. Consignees are to inspect their product and identify if installed units possess the affected software versions. Consignees are to reply to the recall notification confirming that they read and understood the Field Safety Notificaiton. RaySearch is to deploy a new version of RayStation/RayPlan to correct the issue, to be released in April 2024. The provided Field Safety Notice is to be maintained with affected units. Consignees with any questions are to contact quality@raysearchlabs.com. | SoftwareDesign | • Inaccurate reporting of SSD data may lead to incorrect actions. • Potential for reported SSD to be too high may cause incorrect actions. • Erroneous SSD data may affect system operation or performance, leading to incorrect actions. | • Take actions to use the Measure Tool to measure the distance in the patient view to ensure the reported SSD value is correct if it is to be used for patient setup or other critical tasks. • Take actions to inform planning staff and all users of the workarounds provided in the Field Safety Notice. • Take actions to inspect your product and identify if installed units possess the affected software versions. • Take actions to reply to the recall notification confirming that you read and understood the Field Safety Notification. • Take actions to contact quality@raysearchlabs.com if you have any questions.
RayStation | RAYSEARCH LABORATORIES AB | Potential for reported SSD to be too high. | Action A Field Safety Notice for Medical Device Correction #133261, dated 3/28/24, was sent to consignees via email. Consignees were provided the instructions to use the Measure Tool to measure the distance in the patient view to make sure the reported SSD value is correct if it is to be used for patient setup or other critical tasks. If the SSD value is found to be incorrect, small changes to the isocenter, beam angles or the affected ROI's geometry can be made until the correct value is achieved. Consignees are to inform planning staff and all users of the workarounds provided in the Field Safety Notice. Consignees are to inspect their product and identify if installed units possess the affected software versions. Consignees are to reply to the recall notification confirming that they read and understood the Field Safety Notificaiton. RaySearch is to deploy a new version of RayStation/RayPlan to correct the issue, to be released in April 2024. The provided Field Safety Notice is to be maintained with affected units. Consignees with any questions are to contact quality@raysearchlabs.com. | SoftwareDesign | • Incorrect SSD reporting actions may lead to inflated performance expectations and potential system malfunctions. • Inaccurate SSD readings actions can cause unreliable data storage and retrieval, affecting system performance and user trust. • Overreported SSD capacity actions may result in inadequate storage allocation, leading to data loss or corruption. | • Use the Measure Tool to measure the distance in the patient view to ensure the reported SSD value is correct if it is to be used for patient setup or other critical tasks as part of the actions. • Make small changes to the isocenter, beam angles or the affected ROI's geometry if the SSD value is found to be incorrect as part of the actions. • Inform planning staff and all users of the workarounds provided in the Field Safety Notice as part of the actions. • Inspect your product and identify if installed units possess the affected software versions as part of the actions. • Reply to the recall notification confirming that you read and understood the Field Safety Notification as part of the actions. • Contact quality@raysearchlabs.com if you have any questions regarding the actions.
RayStation | RAYSEARCH LABORATORIES AB | Potential for reported SSD to be too high. | Action A Field Safety Notice for Medical Device Correction #133261, dated 3/28/24, was sent to consignees via email. Consignees were provided the instructions to use the Measure Tool to measure the distance in the patient view to make sure the reported SSD value is correct if it is to be used for patient setup or other critical tasks. If the SSD value is found to be incorrect, small changes to the isocenter, beam angles or the affected ROI's geometry can be made until the correct value is achieved. Consignees are to inform planning staff and all users of the workarounds provided in the Field Safety Notice. Consignees are to inspect their product and identify if installed units possess the affected software versions. Consignees are to reply to the recall notification confirming that they read and understood the Field Safety Notificaiton. RaySearch is to deploy a new version of RayStation/RayPlan to correct the issue, to be released in April 2024. The provided Field Safety Notice is to be maintained with affected units. Consignees with any questions are to contact quality@raysearchlabs.com. | SoftwareDesign | • Inaccurate SSD reporting actions may lead to incorrect diagnoses or maintenance. • Potential overestimation of SSD capacity may cause system performance issues or data loss actions. | • Take actions to use the Measure Tool to measure the distance in the patient view to ensure the reported SSD value is correct if it is to be used for patient setup or other critical tasks. • Take actions to inform planning staff and all users of the workarounds provided in the Field Safety Notice. • Take actions to inspect your product and identify if installed units possess the affected software versions. • Take actions to reply to the recall notification confirming that you read and understood the Field Safety Notification. • Take actions to contact quality@raysearchlabs.com if you have any questions.
RayStation | RAYSEARCH LABORATORIES AB | Potential for reported SSD to be too high. | Action A Field Safety Notice for Medical Device Correction #133261, dated 3/28/24, was sent to consignees via email. Consignees were provided the instructions to use the Measure Tool to measure the distance in the patient view to make sure the reported SSD value is correct if it is to be used for patient setup or other critical tasks. If the SSD value is found to be incorrect, small changes to the isocenter, beam angles or the affected ROI's geometry can be made until the correct value is achieved. Consignees are to inform planning staff and all users of the workarounds provided in the Field Safety Notice. Consignees are to inspect their product and identify if installed units possess the affected software versions. Consignees are to reply to the recall notification confirming that they read and understood the Field Safety Notificaiton. RaySearch is to deploy a new version of RayStation/RayPlan to correct the issue, to be released in April 2024. The provided Field Safety Notice is to be maintained with affected units. Consignees with any questions are to contact quality@raysearchlabs.com. | SoftwareDesign | • Inaccurate SSD reporting may lead to incorrect capacity actions. • Potential for reported SSD to be higher than actual capacity, affecting system performance and data storage actions. • Incorrect SSD capacity reporting may cause issues with data storage and system reliability, requiring corrective actions. | • Take actions to measure the distance in the patient view using the Measure Tool to ensure the reported SSD value is correct for patient setup or other critical tasks. • Take actions to make small changes to the isocenter, beam angles, or affected ROI's geometry if the SSD value is found to be incorrect until the correct value is achieved. • Inform planning staff and all users of the workarounds provided in the Field Safety Notice as part of your actions. • Take actions to inspect your product and identify if installed units possess the affected software versions. • Take actions to reply to the recall notification confirming that you read and understood the Field Safety Notification. • Contact quality@raysearchlabs.com if you have any questions or need further actions.
RayStation | RAYSEARCH LABORATORIES AB | Potential for reported SSD to be too high. | Action A Field Safety Notice for Medical Device Correction #133261, dated 3/28/24, was sent to consignees via email. Consignees were provided the instructions to use the Measure Tool to measure the distance in the patient view to make sure the reported SSD value is correct if it is to be used for patient setup or other critical tasks. If the SSD value is found to be incorrect, small changes to the isocenter, beam angles or the affected ROI's geometry can be made until the correct value is achieved. Consignees are to inform planning staff and all users of the workarounds provided in the Field Safety Notice. Consignees are to inspect their product and identify if installed units possess the affected software versions. Consignees are to reply to the recall notification confirming that they read and understood the Field Safety Notificaiton. RaySearch is to deploy a new version of RayStation/RayPlan to correct the issue, to be released in April 2024. The provided Field Safety Notice is to be maintained with affected units. Consignees with any questions are to contact quality@raysearchlabs.com. | SoftwareDesign | • Inaccurate SSD reporting actions may lead to incorrect diagnoses or maintenance. • Overreported SSD capacity may result in unnecessary actions to replace or upgrade. • Potential for incorrect system performance assessments due to inaccurate SSD reporting actions. | • Inspect your product and identify if installed units possess the affected software versions as actions. • Use the Measure Tool to measure the distance in the patient view to ensure the reported SSD value is correct if it is to be used for patient setup or other critical tasks as actions. • If the SSD value is found to be incorrect, make small changes to the isocenter, beam angles or the affected ROI's geometry until the correct value is achieved as actions. • Inform planning staff and all users of the workarounds provided in the Field Safety Notice as actions. • Reply to the recall notification confirming that you read and understood the Field Safety Notice as actions. • Contact quality@raysearchlabs.com if you have any questions as actions.
RayStation | RAYSEARCH LABORATORIES AB | Potential for reported SSD to be too high. | Action A Field Safety Notice for Medical Device Correction #133261, dated 3/28/24, was sent to consignees via email. Consignees were provided the instructions to use the Measure Tool to measure the distance in the patient view to make sure the reported SSD value is correct if it is to be used for patient setup or other critical tasks. If the SSD value is found to be incorrect, small changes to the isocenter, beam angles or the affected ROI's geometry can be made until the correct value is achieved. Consignees are to inform planning staff and all users of the workarounds provided in the Field Safety Notice. Consignees are to inspect their product and identify if installed units possess the affected software versions. Consignees are to reply to the recall notification confirming that they read and understood the Field Safety Notificaiton. RaySearch is to deploy a new version of RayStation/RayPlan to correct the issue, to be released in April 2024. The provided Field Safety Notice is to be maintained with affected units. Consignees with any questions are to contact quality@raysearchlabs.com. | SoftwareDesign | • Incorrect SSD reporting actions may lead to inaccurate system performance assessment. • Potential overestimation of SSD capacity actions can result in reduced system reliability. | • Take actions to measure the distance in the patient view to ensure the reported SSD value is correct using the Measure Tool. • Take actions to make small changes to the isocenter, beam angles, or the affected ROI's geometry if the SSD value is incorrect. • Take actions to inform planning staff and all users of the workarounds provided in the Field Safety Notice. • Take actions to inspect your product and identify if installed units possess the affected software versions. • Take actions to reply to the recall notification confirming that you read and understood the Field Safety Notification. • Contact quality@raysearchlabs.com if you have any questions or need further actions.
RayStation | RAYSEARCH LABORATORIES AB | Potential for reported SSD to be too high. | Action A Field Safety Notice for Medical Device Correction #133261, dated 3/28/24, was sent to consignees via email. Consignees were provided the instructions to use the Measure Tool to measure the distance in the patient view to make sure the reported SSD value is correct if it is to be used for patient setup or other critical tasks. If the SSD value is found to be incorrect, small changes to the isocenter, beam angles or the affected ROI's geometry can be made until the correct value is achieved. Consignees are to inform planning staff and all users of the workarounds provided in the Field Safety Notice. Consignees are to inspect their product and identify if installed units possess the affected software versions. Consignees are to reply to the recall notification confirming that they read and understood the Field Safety Notificaiton. RaySearch is to deploy a new version of RayStation/RayPlan to correct the issue, to be released in April 2024. The provided Field Safety Notice is to be maintained with affected units. Consignees with any questions are to contact quality@raysearchlabs.com. | SoftwareDesign | • Incorrect SSD reporting actions may lead to inaccurate performance metrics. • Overestimated SSD capabilities actions can result in system performance issues. • Potential for reported SSD to be too high actions may cause incorrect system configuration. | • Take actions to inspect your product and identify if installed units possess the affected software versions. • Take actions to inform planning staff and all users of the workarounds provided in the Field Safety Notice. • Take actions to reply to the recall notification confirming that you read and understood the Field Safety Notification. • Take actions to maintain the provided Field Safety Notice with affected units. • Take actions to contact quality@raysearchlabs.com if you have any questions.
Ivenix Infusion System (IIS), LVP Software | Fresenius Kabi USA, LLC | Retroactively reported correction from 2022: A software defect may cause an incorrect (Fail-Stop) Alarm when an administration set is loaded with an upstream occlusion present, resulting in a non-recoverable (Fail-Stop) Pump Problem alarm. This may lead to delay of therapy, and patient may experience from no to minimal to serious harm.******This issue was resolved in all fielded product in software version 5.2.1, which was installed in September 2022. | Action The issue was resolved through an update of the software to Version 5.2.1, which was installed at customer sites in September 2022. The firm did not issue a customer letter at the time. According to the firm, this issue has been resolved in all fielded product. ******There have been two subsequent software upgrades since Version 5.2.1 was released. Consignees are now on v 5.8.0. | SoftwareDesign | • Software defect causes incorrect Fail-Stop Alarm when an administration set is loaded with an upstream occlusion present, leading to non-recoverable Pump Problem alarm actions. • Delay of therapy and potential harm to patients, ranging from minimal to serious, due to the faulty alarm actions. • Potential failure to provide timely treatment, putting patients at risk of adverse health consequences actions. | • Take the actions to ensure the software is updated to Version 5.2.1 or later. • Perform the actions to upgrade to the latest version, which is currently v 5.8.0. • Confirm the actions taken to resolve the issue in all fielded products.
Ivenix Infusion System (IIS), LVP Software | Fresenius Kabi USA, LLC | Retroactively reported corrections from 2023: 1) A software defect may cause an incorrect (Fail-Stop) alarm when an administration set is loaded or coupled while the pump is executing the power-up sequence. May lead to delay in therapy. 2) Alert is not annunciated informing the clinician that the bolus cannot be delivered when the entered bolus dose exceeds the Care Profile Hard Rate Max limit and Rapid Bolus is selected. May lead to over infusion. Both issues were resolved in all fielded product in software version 5.8.0, which was installed in affected units May thru August 2023. | Action Both issues were resolved through an update of the software to Version 5.8.0, which was installed at customer sites between May and August 2023. The firm did not issue a customer letter at the time. According to the firm, this issue has been resolved in all fielded product. | SoftwareDesign | • Incorrect (Fail-Stop) alarm actions may cause delay in therapy when an administration set is loaded or coupled during the pump's power-up sequence. • Failure to annunciate an alert actions may lead to over infusion when the entered bolus dose exceeds the Care Profile Hard Rate Max limit and Rapid Bolus is selected. | • Take the actions to update the software to the latest version, which resolves both issues. • Ensure the actions to install the updated software have been completed.
Synapse PACS | FUJIFILM Healthcare Americas Corporation | Measurements on a Secondary Capture 2D image, that does not have pixel spacing in the DICOM header, when combined with a Breast Tomo Series results in incorrect measurements. | Action URGENT: MEDICAL DEVICE CORRECTION letters were sent to customers January 26, 2024. In an abundance of caution, FUJIFILM is recommending that customer/end users take one of the following actions until the correction is applied: 1. To not perform measurements on Secondary Capture Images that are displayed together with Breast Tomo Images. 2. Ensure Secondary Capture Image has valid pixel spacing values using the DICOM Header Tool. Please complete and return the Field Action Verification Form attached on page 3 of this communication. Providing the response with the information requested is essential for ensuring appropriate action is taken. Please contact our field action coordinator, Kotei Aoki, via phone at 765-246-2931 or email at kotei.aoki@fujifilm.com, if you have any further questions regarding this field action. Please keep a copy of this letter with your instructions for use. | SoftwareDesign | • Incorrect measurements occur when combining a Secondary Capture 2D image without pixel spacing in the DICOM header with a Breast Tomo Series, leading to inaccurate actions. • Failure to include pixel spacing in the DICOM header results in faulty image processing and compromised measurement accuracy, affecting further actions. • Incomplete image header information causes system malfunction, impacting the reliability of subsequent actions. | • Take one of the following actions: do not perform measurements on Secondary Capture Images that are displayed together with Breast Tomo Images or ensure Secondary Capture Image has valid pixel spacing values using the DICOM Header Tool. • Complete and return the Field Action Verification Form to ensure appropriate actions are taken.
Synapse PACS | FUJIFILM Healthcare Americas Corporation | Measurements on a Secondary Capture 2D image, that does not have pixel spacing in the DICOM header, when combined with a Breast Tomo Series results in incorrect measurements. | Action URGENT: MEDICAL DEVICE CORRECTION letters were sent to customers January 26, 2024. In an abundance of caution, FUJIFILM is recommending that customer/end users take one of the following actions until the correction is applied: 1. To not perform measurements on Secondary Capture Images that are displayed together with Breast Tomo Images. 2. Ensure Secondary Capture Image has valid pixel spacing values using the DICOM Header Tool. Please complete and return the Field Action Verification Form attached on page 3 of this communication. Providing the response with the information requested is essential for ensuring appropriate action is taken. Please contact our field action coordinator, Kotei Aoki, via phone at 765-246-2931 or email at kotei.aoki@fujifilm.com, if you have any further questions regarding this field action. Please keep a copy of this letter with your instructions for use. | SoftwareDesign | • Incorrect measurements are performed on a Secondary Capture 2D image due to missing pixel spacing in the DICOM header, leading to inaccurate actions. • Combining this image with a Breast Tomo Series results in faulty actions that compromise the system's performance and reliability. | • Take one of the following actions: do not perform measurements on Secondary Capture Images that are displayed together with Breast Tomo Images, or ensure Secondary Capture Image has valid pixel spacing values using the DICOM Header Tool. • Complete and return the Field Action Verification Form, providing the requested information to ensure appropriate actions are taken.
Synapse PACS | FUJIFILM Healthcare Americas Corporation | Measurements on a Secondary Capture 2D image, that does not have pixel spacing in the DICOM header, when combined with a Breast Tomo Series results in incorrect measurements. | Action URGENT: MEDICAL DEVICE CORRECTION letters were sent to customers January 26, 2024. In an abundance of caution, FUJIFILM is recommending that customer/end users take one of the following actions until the correction is applied: 1. To not perform measurements on Secondary Capture Images that are displayed together with Breast Tomo Images. 2. Ensure Secondary Capture Image has valid pixel spacing values using the DICOM Header Tool. Please complete and return the Field Action Verification Form attached on page 3 of this communication. Providing the response with the information requested is essential for ensuring appropriate action is taken. Please contact our field action coordinator, Kotei Aoki, via phone at 765-246-2931 or email at kotei.aoki@fujifilm.com, if you have any further questions regarding this field action. Please keep a copy of this letter with your instructions for use. | SoftwareDesign | • Incorrect measurements occur when combining a Secondary Capture 2D image without pixel spacing in the DICOM header with a Breast Tomo Series, affecting medical diagnosis and treatment actions. • Lack of pixel spacing in the DICOM header leads to faulty image processing, compromising image quality and accuracy, and potentially impacting patient care actions. | • To prevent any issues, take one of the following actions: do not perform measurements on Secondary Capture Images that are displayed together with Breast Tomo Images. • Take the actions to ensure Secondary Capture Image has valid pixel spacing values using the DICOM Header Tool. • Complete and return the Field Action Verification Form to confirm the actions have been taken. • Contact our field action coordinator if you have any further questions regarding these actions.
Synapse PACS | FUJIFILM Healthcare Americas Corporation | Measurements on a Secondary Capture 2D image, that does not have pixel spacing in the DICOM header, when combined with a Breast Tomo Series results in incorrect measurements. | Action URGENT: MEDICAL DEVICE CORRECTION letters were sent to customers January 26, 2024. In an abundance of caution, FUJIFILM is recommending that customer/end users take one of the following actions until the correction is applied: 1. To not perform measurements on Secondary Capture Images that are displayed together with Breast Tomo Images. 2. Ensure Secondary Capture Image has valid pixel spacing values using the DICOM Header Tool. Please complete and return the Field Action Verification Form attached on page 3 of this communication. Providing the response with the information requested is essential for ensuring appropriate action is taken. Please contact our field action coordinator, Kotei Aoki, via phone at 765-246-2931 or email at kotei.aoki@fujifilm.com, if you have any further questions regarding this field action. Please keep a copy of this letter with your instructions for use. | SoftwareDesign | • Measurements on a Secondary Capture 2D image without pixel spacing in the DICOM header lead to incorrect measurements when combined with a Breast Tomo Series, affecting system performance and accuracy of actions. • Incomplete DICOM header information causes incorrect measurements, compromising data reliability and trust in actions. • System failure to provide accurate measurements impacts user confidence and trust in actions. | • Take one of the following actions: do not perform measurements on Secondary Capture Images that are displayed together with Breast Tomo Images, or ensure Secondary Capture Image has valid pixel spacing values using the DICOM Header Tool, until the correction is applied. • Complete and return the Field Action Verification Form to ensure appropriate actions are taken. • Contact our field action coordinator if you have any further questions regarding this field action.
Synapse PACS | FUJIFILM Healthcare Americas Corporation | Measurements on a Secondary Capture 2D image, that does not have pixel spacing in the DICOM header, when combined with a Breast Tomo Series results in incorrect measurements. | Action URGENT: MEDICAL DEVICE CORRECTION letters were sent to customers January 26, 2024. In an abundance of caution, FUJIFILM is recommending that customer/end users take one of the following actions until the correction is applied: 1. To not perform measurements on Secondary Capture Images that are displayed together with Breast Tomo Images. 2. Ensure Secondary Capture Image has valid pixel spacing values using the DICOM Header Tool. Please complete and return the Field Action Verification Form attached on page 3 of this communication. Providing the response with the information requested is essential for ensuring appropriate action is taken. Please contact our field action coordinator, Kotei Aoki, via phone at 765-246-2931 or email at kotei.aoki@fujifilm.com, if you have any further questions regarding this field action. Please keep a copy of this letter with your instructions for use. | SoftwareDesign | • Incorrect measurements occur when combining a Secondary Capture 2D image without pixel spacing in the DICOM header with a Breast Tomo Series, affecting accurate diagnosis and treatment actions. • Incomplete DICOM header information leads to flawed measurement calculations, compromising patient safety and care quality actions. • System performance is impacted by the inability to accurately process and combine images from different sources, hindering effective diagnosis and treatment actions. | • To take one of the following actions: do not perform measurements on Secondary Capture Images that are displayed together with Breast Tomo Images or ensure Secondary Capture Image has valid pixel spacing values using the DICOM Header Tool, until the correction is applied. • Take actions by completing and returning the Field Action Verification Form. • Contact our field action coordinator if you have any further questions regarding these actions.
Synapse PACS | FUJIFILM Healthcare Americas Corporation | Measurements on a Secondary Capture 2D image, that does not have pixel spacing in the DICOM header, when combined with a Breast Tomo Series results in incorrect measurements. | Action URGENT: MEDICAL DEVICE CORRECTION letters were sent to customers January 26, 2024. In an abundance of caution, FUJIFILM is recommending that customer/end users take one of the following actions until the correction is applied: 1. To not perform measurements on Secondary Capture Images that are displayed together with Breast Tomo Images. 2. Ensure Secondary Capture Image has valid pixel spacing values using the DICOM Header Tool. Please complete and return the Field Action Verification Form attached on page 3 of this communication. Providing the response with the information requested is essential for ensuring appropriate action is taken. Please contact our field action coordinator, Kotei Aoki, via phone at 765-246-2931 or email at kotei.aoki@fujifilm.com, if you have any further questions regarding this field action. Please keep a copy of this letter with your instructions for use. | SoftwareDesign | • Inaccurate measurements result from combining a Secondary Capture 2D image without pixel spacing in the DICOM header with a Breast Tomo Series, requiring corrective actions. • Incorrect data leads to potential misdiagnosis or incorrect treatment, necessitating urgent actions. • System malfunction affects the reliability of medical imaging, demanding immediate corrective actions. | • Take one of the following actions: Do not perform measurements on Secondary Capture Images that are displayed together with Breast Tomo Images. • Take actions to ensure Secondary Capture Image has valid pixel spacing values using the DICOM Header Tool. • Complete and return the Field Action Verification Form to ensure actions are taken. • Contact our field action coordinator for any questions regarding these actions.
Automated Impella Controller (AIC) | Abiomed, Inc. | Retrospective reporting for the release of Technical Bulletin IMP-2643 AIC Version 8.5 Software Update Available. The update resolved an issue in version 8.4 where a pump was not recognized by the AIC. | Action This is a retrospective report for the March 2021 Technical Bulletin, which described a Product Update; specifically the AIC Version 8.5 Software Update. The software update began rollout on March 4, 2021. According to the firm, Version 8.5 of the software contained bug fixes for the issues seen in version 8.4. The technical bulletin was posted to HeartRecovery.com on March 5, 2021. | SoftwareDesign | • The AIC failed to recognize a pump, affecting system operation and performance through incorrect actions. • The failure to recognize the pump resulted in potential safety issues or incorrect actions. | • Take actions to review the retrospective report for the Technical Bulletin, which described a Product Update; specifically the AIC Version 8.5 Software Update. • Take actions to ensure that the software update contains bug fixes for the issues seen in version 8.4.
Synapse PACS | FUJIFILM Healthcare Americas Corporation | The LV Mass (2D Bullet) equation may be calculated incorrectly, resulting in a variance in the mass of the left ventricle (LV Mass). If used as a main factor for diagnosis, there is a risk of misdiagnosis or incorrect treatment plan of a patient, resulting in long-term health consequences or serious deterioration of health. | Action On January 3, 2024, the firm notified customers via URGENT MEDICAL DEVICE CORRECTION letters.******Fujifilm will offer to correct the formula that calculates the LV Mass (2D Bullet) in the Synapse database. Until the update is scheduled and completed, customers may continue to safely use the Pediatric Workflow modules listed in the notice as long as the LV Mass (2D Bullet) equation is not the main factor for a diagnosis.******Customers should coordinate with the Fujifilm service engineer to schedule the correction of their device when contacted.******Please contact Fujifilm's field action coordinator, Kotei Aoki, via phone at 765-246-2931 or email at kotei.aoki@fujifilm.com, if you have any further questions regarding this field action. | SoftwareDesign | • Incorrect calculation of LV Mass (2D Bullet) equation may lead to incorrect diagnosis and treatment actions. • Misdiagnosis or incorrect treatment plan may result in long-term health consequences or serious deterioration of health actions. | • Coordinate with the Fujifilm service engineer to schedule the correction of your device as part of the necessary actions. • Contact Fujifilm's field action coordinator, Kotei Aoki, via phone or email if you have any further questions regarding these actions.
ProxiDiagnost N90 | Philips North America | Issue 1: Mixed Images: The Radio Fluoroscopy (RF) viewer may also display a previous patients radiography (RAD) images when starting the next patient scan while the previous patient image export is still processing. If the issue occurs, there is a potential for incorrect patient data to be displayed in the image.***Issue 2: Detector Access point: Security vulnerability specific to the Wireless Portable Detector configuration items in Philips Support Connect (PSC). Due to this vulnerability it is possible, with physical access to the system and knowledge of specific settings, to modify and export data to removable media (example: USB) | Action ***UPDATE******19 ProxiDiagnost N90 Systems (18 US & 1 OUS) added to scope as these systems were produced prior to implementation of the solution in forward production. The newly identified consignees notified on 20 Nov 2024*********Philips issued URGENT Medical Device Correction ProxiDiagnost N90 1.1 and ProxiDiagnost Upgrade on 12/15/2024. Letter states reason for recall, health risk and action to take:***For Issues 1 and 2:***" To avoid previous patient data from being represented as current patient data, ensure all images ***have been exported from the previous RAD exam before starting a new examination.***" Correct images will be square shape and wrong images can be recognized by a different patient ***anatomy and rectangular shape. If you identify a wrong image, flag the incorrect***images and only export not flagged images for the current examination to PACS (See figure. Refer to Appendix A for specific details regarding the other issue descriptions and advice on actions to be taken.***" Customers can continue using affected systems in accordance with the intended use.***" Circulate this notice to all users of this device so that they are aware of the issue.***Please complete and return the attached response form to Philips promptly and no later***than 30 days from receipt. Completing this form confirms receipt of the URGENT Medical***Device Correction Notice, understanding of the issue, and required actions to be taken.***A Field Service Engineer (FSE) will contact you to schedule a time to visit your site and install the***software upgrade to resolve the issues (Reference FCO70600110). Philips plans to start implementing the solution in Q2 2024.***The software upgrade will also include software enhancements to improve clinical experience.***If you need any further information or support concerning this issue, please contact your local Philips representative: For North America, contact the Customer Care Solutions Center (1-800-722-9377). | SoftwareDesign | • Incorrect patient data may be displayed in the image due to mixed images actions. • Unauthorized data modification and export are possible due to a security vulnerability in the Detector Access point actions. | • To avoid previous patient data from being represented as current patient data, take actions by ensuring all images have been exported from the previous RAD exam before starting a new examination, and only export not flagged images for the current examination to PACS. • Take actions to flag the incorrect images by recognizing them with a different patient anatomy and rectangular shape. • Circulate this notice to all users of this device to make them aware of the issue and the required actions to be taken. • Take actions by completing and returning the attached response form promptly, which confirms receipt of the URGENT Medical Device Correction Notice, understanding of the issue, and required actions to be taken.
ProxiDiagnost N90 | Philips North America | Issue 1: Mixed Images: The Radio Fluoroscopy (RF) viewer may also display a previous patient s radiography (RAD) images when starting the next patient scan while the previous patient image export is still processing. If the issue occurs, there is a potential for incorrect patient data to be displayed in the image.***Issue 2: Detector Access point: Security vulnerability specific to the Wireless Portable Detector configuration items in Philips Support Connect (PSC). Due to this vulnerability it is possible, with physical access to the system and knowledge of specific settings, to modify and export data to removable media (example: USB) | Action ***UPDATE******19 ProxiDiagnost N90 Systems (18 US & 1 OUS) added to scope as these systems were produced prior to implementation of the solution in forward production. The newly identified consignees notified on 20 Nov 2024*********Philips issued URGENT Medical Device Correction ProxiDiagnost N90 1.1 and ProxiDiagnost Upgrade on 12/15/2024. Letter states reason for recall, health risk and action to take:***For Issues 1 and 2:***" To avoid previous patient data from being represented as current patient data, ensure all images ***have been exported from the previous RAD exam before starting a new examination.***" Correct images will be square shape and wrong images can be recognized by a different patient ***anatomy and rectangular shape. If you identify a wrong image, flag the incorrect***images and only export not flagged images for the current examination to PACS (See figure. Refer to Appendix A for specific details regarding the other issue descriptions and advice on actions to be taken.***" Customers can continue using affected systems in accordance with the intended use.***" Circulate this notice to all users of this device so that they are aware of the issue.***Please complete and return the attached response form to Philips promptly and no later***than 30 days from receipt. Completing this form confirms receipt of the URGENT Medical***Device Correction Notice, understanding of the issue, and required actions to be taken.***A Field Service Engineer (FSE) will contact you to schedule a time to visit your site and install the***software upgrade to resolve the issues (Reference FCO70600110). Philips plans to start implementing the solution in Q2 2024.***The software upgrade will also include software enhancements to improve clinical experience.***If you need any further information or support concerning this issue, please contact your local Philips representative: For North America, contact the Customer Care Solutions Center (1-800-722-9377). | SoftwareDesign | • Inaccurate patient data display actions due to mixed images from previous patient scans being displayed during the next patient's scan. • Unauthorized data modification and export actions possible due to a security vulnerability in the Wireless Portable Detector configuration. | • To avoid previous patient data from being represented as current patient data, take the actions to ensure all images have been exported from the previous RAD exam before starting a new examination, and correct images will be square shape and wrong images can be recognized by a different patient anatomy and rectangular shape, and flag the incorrect images and only export not flagged images for the current examination to PACS. • Take the actions to circulate this notice to all users of this device so that they are aware of the issue. • Complete and return the attached response form to Philips, confirming receipt of the URGENT Medical Device Correction Notice, understanding of the issue, and required actions to be taken. • Take the actions to cooperate with the Field Service Engineer (FSE) to schedule a time to visit your site and install the software upgrade to resolve the issues.
HeartMate Touch Communication System | Thoratec Corp. | Due to software and controller systems that results in unexpected pump start or pump stop. | Action On 01/03/2024, the firm sent an Urgent Medical Device Correction Letter via hand delivery to U.S. customers informing them that Abbott has received complaints for the HeartMate Touch Communication System application (App) version 1.0.32 due to unintentional start/stop of the HeartMate 3 Left Ventricle Assist Device (LVAD). ******-Customer should be able to identify the software version of the HeartMate Touch App by tapping "Menu" and "About" on the left-hand panel on the HeartMate Touch screen. ***-When the "STOP PUMP" sequence needs to be performed, follow the instructions in the Instructions For Use (IFU), Chapter 4 HeartMate Touch Communication System Pages 4-58 to 4-59. ***-The user can prevent the unintentional LVAD start/stop from occurring by following the IFU and these additional guidance and recommendations. Note that communication may also be interrupted if the Wireless Adaptor is disconnected from the HeartMate Touch system prior to completion of the STOP PUMP sequence. ***-Do not disconnect the HeartMate System Controller white cable or the Wireless Adapter from the Power Module until STOP PUMP sequence screen with the red progress bar is no longer visible and the STOP PUMP screen changes automatically. ***-During pump priming, do not disconnect the pump until the timer reaches zero, the pump stops, and the Priming is complete message appears as instructed per the IFU.******Abbott is developing and will release a software update to prevent the sequence of events described above. This product correction will be included in an upcoming release of the HeartMate Touch software. An Abbott Representative will contact you for a software upgrade beginning Q2 2024.******For questions, contact Abbott Technical Support at 1-800-456-1477. | SoftwareDesign | • Unexpected pump start or stop actions due to software and controller systems failure. | • You should identify the software version of the HeartMate Touch App by tapping "Menu" and "About" on the left-hand panel on the HeartMate Touch screen, as part of the necessary actions. • When the "STOP PUMP" sequence needs to be performed, follow the instructions in the Instructions For Use (IFU), Chapter 4 HeartMate Touch Communication System Pages 4-58 to 4-59, as part of the required actions. • You should prevent the unintentional LVAD start/stop from occurring by following the IFU and these additional guidance and recommendations, as part of the necessary actions. • Do not disconnect the HeartMate System Controller white cable or the Wireless Adapter from the Power Module until STOP PUMP sequence screen with the red progress bar is no longer visible and the STOP PUMP screen changes automatically, as part of the important actions. • During pump priming, do not disconnect the pump until the timer reaches zero, the pump stops, and the Priming is complete message appears as instructed per the IFU, as part of the necessary actions.
Patient Information Center iX | Philips North America | Push notifications may fail to send to the user under certain conditions. This could potentially result in patient harm due to delay in detection of a change in patient condition. | Action On November 28, 2023, the firm notified affected consignees through URGENT Medical Device Correction letters. This failure is applicable to Apple iOS devices only.******If customers use iOS devices (e.g. iPhone, iPad, etc.) to receive event notifications:******1) Please conduct patient monitoring at the central station or patient bedside.******2) Please do not use iOS event notifications for patient monitoring.******A Philips representative will contact you to schedule a visit to install a software patch to your device***which will correct the current issue. Patch PIC iX 4.2.2 will be provided for PIC iX 4.X and patch CareEvent***C.03.07 will be provided for CareEvent C.03.X.******If you need any further information, please contact your local Philips representative, or call 1-800-722-***9377. | SoftwareDesign | • Failure to send push notifications to users under certain conditions may result in delayed detection of patient condition changes, potentially leading to harmful actions. • Delays in push notification delivery can cause patient harm due to ineffective actions in response to changes in patient condition. | • Take the following actions: Conduct patient monitoring at the central station or patient bedside. • Take the following actions: Do not use iOS event notifications for patient monitoring. • Take the following actions: Wait for a Philips representative to contact you to schedule a visit to install a software patch to your device.
Patient Information Center iX | Philips North America | Push notifications may fail to send to the user under certain conditions. This could potentially result in patient harm due to delay in detection of a change in patient condition. | Action On November 28, 2023, the firm notified affected consignees through URGENT Medical Device Correction letters. This failure is applicable to Apple iOS devices only.******If customers use iOS devices (e.g. iPhone, iPad, etc.) to receive event notifications:******1) Please conduct patient monitoring at the central station or patient bedside.******2) Please do not use iOS event notifications for patient monitoring.******A Philips representative will contact you to schedule a visit to install a software patch to your device***which will correct the current issue. Patch PIC iX 4.2.2 will be provided for PIC iX 4.X and patch CareEvent***C.03.07 will be provided for CareEvent C.03.X.******If you need any further information, please contact your local Philips representative, or call 1-800-722-***9377. | SoftwareDesign | • Failure of push notifications to send to users under certain conditions may delay detection of changes in patient condition, potentially leading to harmful actions. •Delayed detection of patient condition changes due to push notification failures may result in harmful actions. | • Take these actions: conduct patient monitoring at the central station or patient bedside. • Take these actions: do not use iOS event notifications for patient monitoring. • Take these actions: wait for a Philips representative to contact you to schedule a visit to install a software patch to your device. • Take these actions: contact your local Philips representative or call 1-800-722-9377 if you need any further information.
GE Healthcare | GE Healthcare | When using Universal Viewer or Centricity Universal Viewer with Workflow Manager (WFM), if a third-party reporting application is used to launch exams, it could display information on a different patient than WFM and the Viewer. | Action GE Healthcare notified consignees on about 09/08/2023 via letter. Consignees were instructed that until the software update is available the device can still be utilized through an updated browser setting or through discontinuing the use of bi-directional reporting. A GE Healthcare representative will contact each consignee to arrange for the correction. Consignees were instructed to complete and return the provided acknowledgement response. | SoftwareDesign | • Inconsistent patient information display actions when using third-party reporting applications to launch exams, leading to potential errors and mismatches between patient data in Universal Viewer, Centricity Universal Viewer, and Workflow Manager. | • To continue using the device, take actions by updating your browser setting or discontinuing the use of bi-directional reporting. • Take actions by completing and returning the provided acknowledgement response. • Wait for a GE Healthcare representative to contact you to arrange for the correction, and take actions accordingly.
GE Healthcare Centricity | GE Healthcare | When using Universal Viewer or Centricity Universal Viewer with Workflow Manager (WFM), if a third-party reporting application is used to launch exams, it could display information on a different patient than WFM and the Viewer. | Action GE Healthcare notified consignees on about 09/08/2023 via letter. Consignees were instructed that until the software update is available the device can still be utilized through an updated browser setting or through discontinuing the use of bi-directional reporting. A GE Healthcare representative will contact each consignee to arrange for the correction. Consignees were instructed to complete and return the provided acknowledgement response. | SoftwareDesign | • Inconsistent patient information is displayed when a third-party reporting application is used to launch exams, resulting in incorrect patient data actions. • The system fails to synchronize patient information between Universal Viewer, Centricity Universal Viewer, WFM, and the third-party reporting application, leading to incorrect patient data actions. • The mismatch in patient information could lead to incorrect diagnoses, treatments, or medical actions. | • Take the following actions: utilize the device through an updated browser setting or discontinue the use of bi-directional reporting until the software update is available. • Complete and return the provided acknowledgement response as instructed. • Await a GE Healthcare representative's contact to arrange for the correction and follow their actions.
GE Healthcare Centricity | GE Healthcare | GE HealthCare has become aware of an issue where images from two different patients may be contained in a single study, when stored in Centricity PACS-IW V3.7.x, Centricity PACS-IW with Universal Viewer V5.0 and Centricity Universal Viewer V6.0 with PACS-IW foundation. | Action GE Healthcare issued an URGENT MEDICAL DEVICE CORRECTION notice to its consignees on 09/14/2023 by letter. The notice explained the problem, risk to the patient, and provided two workaround options for the continued use of the instruments until the software update is ready for installation. ******Questions: please contact GE HealthCare Service at 1-800-437-1171 or your local Service Representative | SoftwareDesign | • Incorrect patient image association actions: images from two different patients may be contained in a single study, leading to potential misdiagnosis or incorrect treatment actions. | • GE Healthcare has issued an URGENT MEDICAL DEVICE CORRECTION notice explaining the problem and risk to the patient, and provided two workaround options for the continued use of the instruments until the software update is ready for installation, which describe the necessary actions. • Take the following actions and contact GE HealthCare Service at 1-800-437-1171 or your local Service Representative for questions.
GE Healthcare | GE Healthcare | GE HealthCare has become aware of an issue where images from two different patients may be contained in a single study, when stored in Centricity PACS-IW V3.7.x, Centricity PACS-IW with Universal Viewer V5.0 and Centricity Universal Viewer V6.0 with PACS-IW foundation. | Action GE Healthcare issued an URGENT MEDICAL DEVICE CORRECTION notice to its consignees on 09/14/2023 by letter. The notice explained the problem, risk to the patient, and provided two workaround options for the continued use of the instruments until the software update is ready for installation. ******Questions: please contact GE HealthCare Service at 1-800-437-1171 or your local Service Representative | SoftwareDesign | • Images from two different patients may be contained in a single study, leading to incorrect medical actions. • Incorrect patient information may lead to incorrect medical actions. | • Take actions to address the problem with GE Healthcare instruments by following two workaround options until the software update is ready for installation. • Contact GE HealthCare Service at 1-800-437-1171 or your local Service Representative for further actions.
Beckman Coulter DxI 9000 Access Immunoassay Analyzer | Beckman Coulter, Inc. | The firm has investigated and confirmed that when DxI 9000 is connected to Laboratory Information System (LIS, host), the sample barcode presented may be ignored and replaced with a sample ID that was previously in the same rack and position. This will cause a delay of result to the sample currently presented to the system (Sample B) and may result in an erroneous test result for the prior sample (Sample A), which gets associated with***the results of sample B. | Action Beckman Coulter notified its consignees on 09/12/2023 via letter. The notice explained the issue, risk, and consignees were advised to perform a retrospective review of all results that were generated on a DxI 9000 analyzer that is connected to a host system (laboratory information system or middleware).***Current customers should perform one of the following:***Option 1: Disconnect the DxI 9000 Access***Immunoassay Analyzer from the host system. You may continue ordering tests through the analyzer console.***Option 2: Prevent the host system from downloading test orders or sending new information updates to the analyzer.***Beckman Coulter will release system software version 1.15.2, which prevents reporting an erroneous test result. Test results will still be delayed when the DxI 9000 Access Immunoassay Analyzer may incorrectly identify samples. | SoftwareDesign | • Sample barcode may be ignored and replaced with a previous sample ID, causing delayed and erroneous test results through incorrect sample association actions. • Erroneous test results may be associated with a prior sample (Sample A) due to the system's malfunction, leading to incorrect actions. • Delayed results may occur for the currently presented sample (Sample B) due to the system's failure to recognize the correct sample barcode, affecting timely actions. | • To take immediate actions, disconnect the DxI 9000 Access Immunoassay Analyzer from the host system, allowing continued ordering of tests through the analyzer console as necessary actions. • As an alternative, take actions to prevent the host system from downloading test orders or sending new information updates to the analyzer.
DxA 5000 (DxA Automation System, DxA 5000 Fit) | Beckman Coulter Biomedical GmbH | A software defect in the DxA Automation System may cause tests to be run on patient samples in tube types not compatible with the requested test, leading to erroneous patient results | Action On 06/12/2023, Beckman Coulter issued an Urgent Medical Device Recall Notification to affected consignees. Beckman Coulter asked consignees to take the following actions:******1. Inspect tubes routed to error regions and confirm that the tube type was appropriate per the assay IFU for the test performed.******2. Please share this information with your laboratory staff and retain this notification as part of your laboratory Quality System documentation. If you have forwarded any of the affected product(s) listed above to another laboratory, please provide them a copy of this letter.******3. So that we are assured you have received this important communication, please respond within 10 days in one of the following ways:***" Electronically, if you received this communication via email./" Manually, complete and return the enclosed Response Form.***4. If you have any questions regarding this notice, please contact our Customer Support Center or your local Beckman Coulter Representative:***" From our website: http://www.beckmancoulter.com/" By phone: call 1-800-526-3821 in the United States.***5. Adverse reactions or quality problems experienced with the use of this product may be reported to the FDA s MedWatch Adverse Event Reporting program either online, by regular mail or by fax. | SoftwareDesign | • Erroneous patient results due to incompatible tube types used for tests, affecting the accuracy of diagnoses and subsequent medical actions. • Incompatible tube types lead to incorrect test results, potentially causing incorrect medical actions or treatment decisions. • Defect in the DxA Automation System may lead to unintended actions, compromising patient safety and care. • Inaccurate test results may prompt incorrect actions, such as prescribing unnecessary medications or performing unnecessary procedures. | • Inspect tubes and confirm appropriate tube type per assay IFU for the test performed as part of the actions. • Share this information with laboratory staff and retain it as part of laboratory Quality System documentation as part of the actions. • Respond within 10 days to confirm receipt of this communication as part of the actions. • Contact our Customer Support Center or local Beckman Coulter Representative if you have questions regarding this notice as part of the actions. • Report adverse reactions or quality problems to the FDA's MedWatch Adverse Event Reporting program as part of the actions.
DxA 5000 (DxA Automation System, DxA 5000 Fit) | Beckman Coulter Biomedical GmbH | A software defect in the DxA Automation System may cause tests to be run on patient samples in tube types not compatible with the requested test, leading to erroneous patient results | Action On 06/12/2023, Beckman Coulter issued an Urgent Medical Device Recall Notification to affected consignees. Beckman Coulter asked consignees to take the following actions:******1. Inspect tubes routed to error regions and confirm that the tube type was appropriate per the assay IFU for the test performed.******2. Please share this information with your laboratory staff and retain this notification as part of your laboratory Quality System documentation. If you have forwarded any of the affected product(s) listed above to another laboratory, please provide them a copy of this letter.******3. So that we are assured you have received this important communication, please respond within 10 days in one of the following ways:***" Electronically, if you received this communication via email./" Manually, complete and return the enclosed Response Form.***4. If you have any questions regarding this notice, please contact our Customer Support Center or your local Beckman Coulter Representative:***" From our website: http://www.beckmancoulter.com/" By phone: call 1-800-526-3821 in the United States.***5. Adverse reactions or quality problems experienced with the use of this product may be reported to the FDA s MedWatch Adverse Event Reporting program either online, by regular mail or by fax. | SoftwareDesign | • Software defect in DxA Automation System causes tests to be run on incompatible tube types, leading to actions that produce erroneous patient results. | • Inspect tubes routed to error regions and confirm that the tube type was appropriate per the assay IFU for the test performed as part of the required actions. • Share this information with your laboratory staff and retain this notification as part of your laboratory Quality System documentation as part of the required actions. • Respond within 10 days in one of the following ways as part of the required actions: electronically or manually, complete and return the enclosed Response Form. • Contact our Customer Support Center or your local Beckman Coulter Representative if you have any questions regarding this notice as part of the required actions. • Report any adverse reactions or quality problems experienced with the use of this product to the FDA's MedWatch Adverse Event Reporting program as part of the required actions.
ARTIS pheno | Siemens Medical Solutions USA, Inc | If, during the procedure, X-ray has been released and a reference image has been stored, the following issue may occur: If "Adjust C-arm to Ref" is activated when the C-Arm is positioned outside of the working range the C-Arm will reach the target position with an inaccuracy of 5-10 mm. As a result, the message Endposition reached will be displayed .Leading to unintended direction of the movement causing crushing of a patient, staff member, operator, or equipmen******Live images may not match the previously stored reference images. Overlay images may be shown inaccurate on anatomy (e.g., DSA Roadmap workflow does not match real anatomy). This may cause e.g., a vessel perforation in DSA roadmap.*** | Action Siemens Healthineers issued to all users of ARTIS pheno systems with a Siemens Healthineers table or Trumpf/MAQUET table (OEM) a Customer Safety Advisory Notice to affected customers via AX041/22/S dated 2/3/23. Letter states reason for recall, health risk and action to take for addressing three potential software issues: Detection of wrong table movement direction; Adjust C-arm to Ref; Missing movement stop with multiple collision sensor activation. Siemens will correct the software error via Update Instructions AX040/22/S. Our service organization will contact you shortly to arrange a date to perform this corrective action. Please feel free to contact our service organization for an earlier appointment at 1-800-888-7436. Please review, understand, and confirm acknowledgement of the information and instructions outlined in this Urgent Safety Advisory Notice.******Siemens ******A software update to correct the error will be distributed to all affected customers via Update Instructions AX040/22/S. Affected systems will be updated to SW version VE10B. | SoftwareDesign | • Inaccurate C-Arm positioning actions due to incorrect reference image, leading to unintended direction of movement and potential crushing of patients, staff, or equipment. • Inaccurate overlay images actions, mismatching real anatomy, which can cause vessel perforation in DSA roadmap workflow. | • Review, understand, and confirm acknowledgement of the information and instructions outlined in this Urgent Safety Advisory Notice, which describes actions to address three potential software issues. • Wait for our service organization to contact you to arrange a date to perform the corrective actions. • Feel free to contact our service organization for an earlier appointment at 1-800-888-7436 to expedite the actions. • Update your system to SW version VE10B following the actions outlined in Update Instructions AX040/22/S to correct the software error.
ARTIS pheno | Siemens Medical Solutions USA, Inc | In the event of any unintended table movement, the system may not detect the incorrect direction, could lead to the injury of a patient, staff member, operator, or equipment. | Action Siemens Healthcare issued A Customer Safety Advisory Notice distributed to customers on 2/2/23 via AX041/22/S. A software update to correct the error will be distributed to all affected customers via Update Instructions AX040/22/S. Affected systems will be updated to SW version VE10B. Letter states reason for recall, health risk and action to take:***Please perform system movements with adequate care to avoid injuries to persons and damage to equipment due to collisions or crushing.******Siemens will correct the software error via Update Instructions AX040/22/S. Our service organization will contact you shortly to arrange a date to perform this corrective action. Please feel free to contact our service organization for an earlier appointment at 1-800-888-7436. Please review, understand, and confirm acknowledgement of the information and instructions outlined in this Urgent Safety Advisory Notice. By signing this document, you are acknowledging that you have read and understand the content. | SoftwareDesign | • Unintended table movement may lead to incorrect direction detection, compromising system operation and putting patients, staff members, operators, or equipment at risk of injury through incorrect actions. • Failure to detect incorrect direction can result in physical harm to individuals or damage to equipment, necessitating correct actions to prevent accidents. | • Please perform system movements with adequate care to avoid injuries to persons and damage to equipment due to collisions or crushing as part of the corrective actions. • Follow the instructions outlined in this Urgent Safety Advisory Notice as part of the necessary actions. • Confirm acknowledgement of the information and instructions by signing this document to acknowledge you have read and understand the content, which is part of the required actions. • Contact our service organization at 1-800-888-7436 to arrange a date to perform the corrective actions or to schedule an earlier appointment.
ExacTrac Dynamic | Brainlab AG | Deep Inspiration Breath Hold (DIBH) functionality may not work as specified when based on a standard-workflow Treatment Template | Action The recalling firm issued letters dated 1/25/2023 via email and were flagged URGENT FIELD SAFETY NOTICE / PRODUCT NOTIFICATION. The letter provides the user information on how the issue occurs and informs them of the corrective actions Brainlab is taking to address the issue. Detailed information on the issue and its effects are provided. The letter provides User Corrective Actions: (1) Always create DIBH Treatment Templates from the Brainlab provided DIBH Master Templates; and (2) Always ensure that Beam Hold Control is enabled ("On") before proceeding to DIBH treatment. Brainlab will provide a software revision of ExacTrac Dynamic to all affected customers with the described issue corrected. Brainlab will actively contact the consignee to schedule the update starting April 2023. The consignee is to advise the appropriate personnel working in their department of the content of the letter. Consignee response to the letter was requested in the email. The response by the consignee was to include they have received the notification, the referenced device is installed at their hospital and they are the correct recipient for this communication, and the users at the hospital have been made aware of the content and to adhere to the User Corrective Action described in the notification. | SoftwareDesign | • DIBH functionality malfunctions when based on a standard-workflow Treatment Template, affecting actions. • The malfunction can lead to inconsistent or incorrect treatment outcomes, compromising actions. • The recall affects the system's operation, compromising patient safety and treatment efficacy, requiring actions. | • Take the following actions: Always create DIBH Treatment Templates from the Brainlab provided DIBH Master Templates; and Always ensure that Beam Hold Control is enabled ("On") before proceeding to DIBH treatment. • Advise the appropriate personnel working in your department of the content of this notification and ensure they adhere to the actions described above.
Oncentra Brachy | Elekta, Inc. | A reconstruction error may occur when using the Catheter Bending functionality in Applicator Modeling or Implant Modeling | Action Elekta issued an Important Field Safety Notice 806-01-BTP-003 dated November 2022 to all affected customers on 12/5/2022 via email. The notice informs users of the specific product and version numbers affected by the issue and any work arounds that can be used to avoid the issue. The notice is to be posted in a place accessible to all users until this action is closed. Appropriate personnel working with the product are to be advised of the content of the letter. The customers are instructed to complete and return the Important Field Safety Notice Acknowledgement form to their local Elekta representative as soon as possible but no longer than 30 days from receipt.******Update: The manufacturer has now released the solution to this issue for affected Oncentra Brachy users using version 4.0 and above with Applicator Modeling license 170283-00. All affected customers will receive a software update in order to solve the issue. The solution was released for installation in the field on 20-MAR-23. Customers are being notified by email on 03/20/2023. Customers are being informed that the fix is available for download/request depending on their standard processes. Elekta service teams will pro-actively follow up with customers to insure the fix has been installed before the target completion date. | SoftwareDesign | • A reconstruction error may occur, resulting in inaccurate modeling actions. • Catheter Bending functionality may malfunction, affecting implant/applicator modeling actions. • Inaccurate modeling may impact system performance and reliability, leading to potential safety issues with medical actions. | • Post the Important Field Safety Notice in a place accessible to all users until this action is closed, ensuring that all users are informed of the affected product and version numbers. • Advise appropriate personnel working with the product of the content of the notice, highlighting the necessary actions. • Complete and return the Important Field Safety Notice Acknowledgement form to the local Elekta representative, indicating that the necessary actions have been taken. • Take actions to apply the software update to resolve the issue, which will be provided by Elekta, and ensure that the fix is installed before the target completion date.
ARTIS pheno | Siemens Medical Solutions USA, Inc | Siemens has become aware of three potential software issues with ARTIS pheno systems with software version VE10B. This may lead to a hazardous situation for patients if treatment cannot be continued on the system and treatment needs to be continued on an alternate system. | Action On September 20, 2022, the firm notified affected customers via Urgent Medical Device Correction letters. This notification covered three issues.******1. System Error Management, in which the operator is unable to release X-Ray in "Bypass fluoroscopy" mode which remains permanently inhibited, and may result in a situation where it is necessary to cancel clinical treatment or to continue treatment on an alternative system. Customers should ensure that patient treatment can be continued in other ways if there is any possible danger for the safety of the patient.******2. Head Holder Symbols, in which the user cannot tell whether a head holder has been selected in the system, may lead to a collision between the C-Arm and the head holder during system movement if the head holder is being used but not selected on the system. Customers should ensure that the configuration of the head holder is correct prior to system movements and double check configuration in the control room in case of uncertainty during the procedure.******3. Stand Movement, in which rarely no stand movement is possible after a routine brake test, leading to a possible delay of procedure. Customers should carry out routine brake tests with sufficient time before starting any procedure.******Siemens will correct the software errors and contact customers to arrange a date to perform the corrective action. Please feel free to contact Siemens' service organization for an earlier appointment at 1-800-888-7436. | SoftwareDesign | • Software issues with ARTIS pheno systems may lead to hazardous situations for patients if treatment cannot be continued, requiring alternative actions. • Potential system failure may disrupt ongoing treatment, necessitating urgent alternative actions. • System unavailability may cause delays or interruptions, requiring immediate alternative actions. | • Ensure patient treatment can be continued in other ways if there is any possible danger for the safety of the patient as part of actions. • Ensure the configuration of the head holder is correct prior to system movements and double check configuration in the control room in case of uncertainty during the procedure as part of actions. • Carry out routine brake tests with sufficient time before starting any procedure as part of actions.
Merlin Patient Care System (PCS) Software | St. Jude Medical, Cardiac Rhythm Management Division | There is a potential for programmer software (Merlin PCS) and remote monitoring software application (Merlin.net) to display overestimated predicted battery longevity for certain pacemakers. | Action On June 16, 2022 Abbott issued a "Medical Device Correction" via: XXXX to affected consignees. The customer notification ask consignees to take the following action: ******1. Prophylactic device replacement is not recommended, as device functionality, actual longevity, and ERI indicator are not impacted (device functionality remains normal and within specifications).******2. Routine follow-up should remain as per local standard of care and clinical protocol, and ERI should continue to serve as an indicator of the need for device replacement scheduling.******3. Please direct any questions about device longevity to Abbott Technical Support at 1-800-722-3774 (U.S.).******4.Upon programmer software / remote monitoring software update, the improved longevity estimate will be displayed at the patient s next interrogation. Please note that until programmers are updated, a difference in longevity estimates between programmers and remote monitoring (Merlin.net) may be observed. | SoftwareDesign | • Displaying overestimated predicted battery longevity for certain pacemakers through programmer software and remote monitoring software application actions. | • Take the following actions: do not replace the device prophylactically as its functionality and longevity are not impacted. • Take the following actions: continue routine follow-up as per local standard of care and clinical protocol, and use the ERI as an indicator for device replacement scheduling. • Take the following actions: direct any questions about device longevity to Abbott Technical Support at 1-800-722-3774. • Take the following actions: expect the improved longevity estimate to be displayed at the patient's next interrogation upon programmer software / remote monitoring software update.
GE Centricity | GE Healthcare, LLC | Potential to display inaccurate measurements on images in Centricity Universal Viewer Zero Footprint Client (ZFP) | Action GE notified their consignees by issuing an URGENT MEDICAL DEVICE CORRECTION notice on 08 June 2022. The letter explains the two safety issues and provides the following actions be taken:***Issue #1: "You can continue to use your device in accordance with the User Manual and the actions below:***Do not utilize any measurements when the 'Lossy' overlay is displayed in the lower left of the viewport and, when you hover over the word 'Lossy', it displays the tooltip box with the phrase 'Image scaled down from its original resolution'."******Issue #2: "You can continue to use your device in accordance with the User Manuals and the actions below:***Do not perform any measurement on non-squared pixel images. Always confirm critical diagnostic measurements by other means (i.e using a fiducial marker of known size)."******GE Healthcare will correct all affected products at no cost to you. A GE Healthcare representative will contact you to arrange for the correction.***After the software has been corrected, be sure to destroy the installation media for affected software at your site.******If you have any questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. | SoftwareDesign | • Inaccurate measurements may be displayed on images in Centricity Universal Viewer Zero Footprint Client (ZFP) due to faulty image rendering actions. • Potential for inaccurate diagnoses or treatments may occur as a result of relying on incorrect measurement data actions. | • Take the following actions: Do not utilize any measurements when the 'Lossy' overlay is displayed in the lower left of the viewport and, when you hover over the word 'Lossy', it displays the tooltip box with the phrase 'Image scaled down from its original resolution'. • Take the following actions: Do not perform any measurement on non-squared pixel images, and always confirm critical diagnostic measurements by other means (i.e., using a fiducial marker of known size). • Take the following actions: After the software has been corrected, be sure to destroy the installation media for affected software at your site.
ACUSON Juniper Diagnostic Ultrasound System | Siemens Medical Solutions USA, Inc. | The clip store function in the ultrasound imaging system does not work when the system has a disk full error. This could cause a delay in treatment if the ultrasound system is unable to save clips as study documentation during a high risk procedure, such as a stress echo exam. | Action The firm, Siemens Healthineers, sent "URGENT MEDICAL DEVICE SAFETY CORRECTION" Letters Juniper 1.0 (VA10D/VA10E/VA10F) to affected customers via Siemens update program US026/21/S informing them of the issue and advising them the firm will correct this issue with a free-of-charge software update to their ACUSON Juniper ultrasound system. The customers were instructed to do the following: Check the available storage on the ultrasound system prior to starting a study where clips are vital to the ***documentation, such as stress echo. If available, use a different ultrasound system capable of performing a stress echo exam until the problem is resolved. If another ultrasound system is not available, consider using other imaging modalities to achieve the diagnostic outcome. In the event that a defect is encountered, a reinstallation of the system software can resolve this problem until the available storage on the ultrasound system is again exceeded.******The customers' respective customer service engineer will contact them to schedule a facility visit to update the system or inform them of a remote update.******If you have further questions regarding this safety correction, please send all inquiries to: Siemens Healthineers USD Complaints usd-complaint.team@siemens-healthineers.com. | SoftwareDesign | • Clip store function doesn't work when the system has a disk full error, affecting actions. • Delay in treatment may occur when clips cannot be saved as study documentation during high-risk procedures, affecting actions. | • To ensure the safe operation of your ACUSON Juniper ultrasound system, take the following actions: check the available storage on the system prior to starting a study where clips are vital to documentation. • Take the following actions: use a different ultrasound system capable of performing a stress echo exam until the problem is resolved, if possible. • If another ultrasound system is not available, take the following actions: consider using other imaging modalities to achieve the diagnostic outcome. • In the event of a defect, take the following actions: reinstallation of the system software can resolve this problem until the available storage on the ultrasound system is again exceeded.
RaySearch | RAYSEARCH LABORATORIES AB | If a new primary image set is selected while the cine loop is running, the primary image set will be displayed as both primary and secondary image set in all side-by-side views. This will also be true for any new patient or case opened while the cine loop is running | Action RaySearch issued Field Safety Notice, Medical Device Correction #90375 via*** E-mail, on 30 December 2021. Letter stated reason for recall, health risk and action to take:***Do not change the primary image set while the cine loop is running.***" If the primary image set has been changed while the cine loop is running, restart***RayStation/RayPlan to restore correct behavior.***" Educate planning staff and all users about this workaround.***" Inspect your product and identify all installed units with the above software version number(s).***" Confirm you have read and understood this notice by replying to the notification email.***Solution.******This issue is resolved in RayStation/RayPlan versions 10A and higher (subject to market clearance in somemarkets). If customers wish to continue using versions of RayStation/RayPlan affected by this notice, all users must maintain awareness of this notice. Alternatively, customers can choose to upgrade to the new version once it becomes available for clinical use.***Please confirm that you have received this FSN.***If you want to attach a signed reply form to the email, please fill in the below. You can also fax this form to Fax: +1-631-828-2137 (US only).******For regulatory information, please contact quality@raysearchlabs.com. | SoftwareDesign | • Inconsistent image display actions occur when a new primary image set is selected during a cine loop, causing the primary image set to be shown as both primary and secondary in side-by-side views. • New patient or case openings during a cine loop also trigger incorrect image display actions, leading to primary image sets being displayed as both primary and secondary. | • Educate planning staff and all users about the actions to take: do not change the primary image set while the cine loop is running, and restart RayStation/RayPlan to restore correct behavior if it has been changed. • Inspect your product and identify all installed units with the affected software version number(s) and take actions accordingly. • Confirm you have read and understood this notice by replying to the notification email and take actions to maintain awareness of this notice if continuing to use affected versions. • Take actions to upgrade to the new version of RayStation/RayPlan once it becomes available for clinical use as an alternative to maintaining awareness of this notice.
ExacTrac Dynamic | Brainlab AG | The yaw angle may be incorrect for CBCT positioning workflows using setup beams with Varian LINACs. | Action The recalling firm initiated the recall via email on 2/25/2022 with the Urgent Field Safety Notice dated 2/7/2022 attached. The email says to ensure the safe and effective use of the device, the consignee is to reply (1) they have received the Field Safety Notice; (2) the affected device is installed at the hospital and the the addressee of the email is the correct recipient for the communication; and (3) users at the hospital have been made aware of the content and adhere to the User Corrective Action described in the enclosed notification. The email asks the consignee to answer each of the 3 points to ensure the effectiveness of the communication.******The notice explained the issue and provided details of specific conditions that are required for the error to occur, the affected workflows, and the magnitude of the potential error. User Corrective Action is provided giving instructions to prevent the issue from occurring. The action provided was: When using the ETD CBCT module, always ensure that during the CBCT X-ray Reference Acquisition step (CBCT offset detection), a treatment beam and not a setup beam is loaded on the Varian treatment console. If a setup beam is loaded, the word Setup will display in the ETD software, and it is not safe to continue. (A figure is provided to show the display. It shows a screen where a setup beam is active and it is not safe to continue and also shows the display of a treatment beam that is active and it is safe to continue.) In this case, the user is to assure they first switch to a treatment beam on the Varian treatment console and confirm the change in the Brainlab ETD software before continuing.******Brainlab will provide a software revision of ExacTrac Dynamic with the described issue corrected to all affected customers. Brainlab will actively contact the consignee to schedule the update, starting in May 2022.******Appropriate personnel working in the consignee's department are to be made aware of the notification. | SoftwareDesign | • Inaccurate yaw angle may lead to incorrect CBCT positioning workflows using setup beams with Varian LINACs, affecting treatment accuracy and requiring corrective actions. • Potential misalignment of setup beams may compromise radiation therapy treatment outcomes, necessitating mitigating actions. | • Ensure safe and effective use of the device by taking the following actions: reply to the notification, confirm the affected device is installed and you are the correct recipient, and ensure users are aware of the content and adhere to the User Corrective Action. • Take the following actions when using the ETD CBCT module: always ensure a treatment beam and not a setup beam is loaded on the Varian treatment console, and switch to a treatment beam and confirm the change in the Brainlab ETD software before continuing. • Ensure all appropriate personnel working in your department are aware of the notification and take necessary actions.
GE Centricity | GE Healthcare, LLC | There is a potential to display incomplete patient imaging study. | Action The firm issued medical device correction letters dated 10/15/2021 on the same date via traceable means. The letter described the safety issue and actions to be taken. The letter informed the consignee they can continue to use the device by following the instructions provided in the letter until a GE Service Representative can restart their ZFP system. | SoftwareDesign | • Incomplete patient imaging studies may be displayed, requiring corrective actions. • Potential for incomplete imaging data may lead to misdiagnosis or inadequate treatment, necessitating corrective actions. • System malfunction may compromise patient safety and care, demanding immediate corrective actions. | • The letter informed the consignee of the necessary actions to continue using the device safely. • The actions include following the instructions provided in the letter until a service representative can restart their system.
HeartMate Touch Communication System | Thoratec Corp. | If LVAS communication system is trying to establish Bluetooth connection and another Bluetooth-enabled device is nearby and advertising for Bluetooth connection, the Bluetooth connectivity interference may cause the system App to close or fail to open. Once Bluetooth adaptor is connected, barring disconnection, all functionality is unaffected; the device will continue to operate as expected. | Action On 10/22/21, correction notices were provided via site visit, or email. ******Until the resolution is provided, the recalling firm is requesting that your facility retain or obtain the System Monitor II (Model Numbers 1286 and L1286) and ensure it is readily available for use in the event that the LVAS communication system is unable to function properly. The recalling firm is extending the support and servicing of System Monitors II until this issue is resolved and the resolution is fully implemented in the field. If your facility does not have a System Monitor II, the recalling firm will provide a loaner System Monitor II at no charge until a corrective action is provided to your facility. If your facility has not previously used a System Monitor II, the recalling firm will provide training and implementation support to your facility. A firm representative will work with you to schedule a time that is convenient for you and your staff.******If you have any questions about this communication or the recommendations, please contact your local representative or Technical Support at 1-800-456-1477 (U.S.). Please share this notice with all those within your organization who might operate the affected device. We ask that you maintain a record of this notice to ensure effectiveness of the communication. Please complete and return the response form. | SoftwareDesign | • Bluetooth connectivity interference actions may cause the system App to close or fail to open when another Bluetooth-enabled device is nearby and advertising for Bluetooth connection. • Failure to establish a stable Bluetooth connection actions may result in disruption to system operation and functionality. | • Retain or obtain the System Monitor II (Model Numbers 1286 and L1286) and ensure it is readily available for use in the event that the LVAS communication system is unable to function properly as instructed by the recalling firm in their actions. • If you don't have a System Monitor II, request a loaner from the recalling firm at no charge until a corrective action is provided to your facility, following their actions. • If you haven't previously used a System Monitor II, request training and implementation support from the recalling firm, as part of their actions. • Contact your local representative or Technical Support at 1-800-456-1477 (U.S.) if you have any questions about this communication or the recommendations, regarding the actions. • Share this notice with all those within your organization who might operate the affected device as part of the actions. • Maintain a record of this notice to ensure effectiveness of the communication, as per the actions. • Complete and return the response form, as instructed in the actions.
Siemens | Siemens Medical Solutions USA, Inc | software syngo.CT VB20 in the installed base, with or without syngo.CT software versions VB20_SP1, SP2, SP3, or SP4, may result in workflow interruptions***resulting in the delay in diagnosis, the necessity for patient rescans, and/or the need for additional contrast media may occur | Action Siemens Medical Solutions USA, Inc. initiated a Customer Safety Advisory Notice (CSAN) to customers via CT021/21/S on 8/30/2021. The CSAN informs the customer of the issues health risk, the new software update version VB20_SP5 to eliminate the root causes of these problems. The software update will be distributed free of charge via CT020/21/S and CT071/21/S (update is dependent on the current software version) to all affected customers. ***Action to take:***Siemens Healthineers has developed software update syngo.CT VB20 SP5. This new software contains workflow improvements and solutions to the previously described software issues. Please ask your local service organization if you are interested in more detailed information about the content of the software update.***The new software, VB20_SP5, will be provided free of charge and will be distributed via Update CT020/21/S or CT071/20/S. Update CT071/20/S is only necessary if the system is being updated from software version VB20, VB20_SP1, or VB20_SP2 directly to VB20_SP5. Following the corrective action process CAPA, the cause of identified***product issues has been eliminated and a recurrence of the identified issues is prevented.***This update will be performed remotely. The remote software updates process will require approximately 45 minutes for completion and is completely automatic; however, please make sure the system and power are stable before and during the process, please do not switch off the system during the update process. Siemens highly recommends***starting the installation when the scanner will not be in use or when the necessary time for the update to be completed can be scheduled.***If you have any questions, please contact our service organization at at 1-800-888-7436.******Please review, understand, and confirm acknowledgement of the information and instructions outlined in this***Urgent Safety Advisory Notice. By signing this document, you are acknowledging that you have read and understand the content therein | SoftwareDesign | • Workflow interruptions actions may result in delayed diagnosis. • Patient rescans actions may be necessary due to software malfunction. • Additional contrast media actions may be required due to software failure. | • Contact your local service organization to obtain more information about the software update syngo.CT VB20 SP5 and its content, which contains workflow improvements and solutions to the previously described software issues, and take actions accordingly. • Ensure a stable system and power supply before and during the remote software update process, and take no actions to switch off the system during the update, which will take approximately 45 minutes to complete. • Take actions to schedule the software update when the scanner will not be in use or when the necessary time for the update to be completed can be scheduled, as highly recommended. • Contact our service organization at 1-800-888-7436 if you have any questions or concerns, and take actions to resolve them promptly. • Take actions to review, understand, and confirm acknowledgement of the information and instructions outlined in this Urgent Safety Advisory Notice.
Siemens | Siemens Medical Solutions USA, Inc | software syngo.CT VB20 in the installed base, with or without syngo.CT software versions VB20_SP1, SP2, SP3, or SP4, may result in workflow interruptions***resulting in the delay in diagnosis, the necessity for patient rescans, and/or the need for additional contrast media may occur | Action Siemens Medical Solutions USA, Inc. initiated a Customer Safety Advisory Notice (CSAN) to customers via CT021/21/S on 8/30/2021. The CSAN informs the customer of the issues health risk, the new software update version VB20_SP5 to eliminate the root causes of these problems. The software update will be distributed free of charge via CT020/21/S and CT071/21/S (update is dependent on the current software version) to all affected customers. ***Action to take:***Siemens Healthineers has developed software update syngo.CT VB20 SP5. This new software contains workflow improvements and solutions to the previously described software issues. Please ask your local service organization if you are interested in more detailed information about the content of the software update.***The new software, VB20_SP5, will be provided free of charge and will be distributed via Update CT020/21/S or CT071/20/S. Update CT071/20/S is only necessary if the system is being updated from software version VB20, VB20_SP1, or VB20_SP2 directly to VB20_SP5. Following the corrective action process CAPA, the cause of identified***product issues has been eliminated and a recurrence of the identified issues is prevented.***This update will be performed remotely. The remote software updates process will require approximately 45 minutes for completion and is completely automatic; however, please make sure the system and power are stable before and during the process, please do not switch off the system during the update process. Siemens highly recommends***starting the installation when the scanner will not be in use or when the necessary time for the update to be completed can be scheduled.***If you have any questions, please contact our service organization at at 1-800-888-7436.******Please review, understand, and confirm acknowledgement of the information and instructions outlined in this***Urgent Safety Advisory Notice. By signing this document, you are acknowledging that you have read and understand the content therein | SoftwareDesign | • Workflow interruptions may result in delayed diagnosis actions. • Patient rescans may be necessary due to software issues actions. • Additional contrast media may be required due to software failures actions. | • Please ask your local service organization about the software update syngo.CT VB20 SP5, which includes workflow improvements and solutions to previously described software issues, and its actions. • Verify if the update is necessary for your system and its actions. • Ensure the system and power are stable before and during the remote software update process, which will take approximately 45 minutes to complete, and do not switch off the system during the process, and its actions. • Schedule the update when the scanner will not be in use or when the necessary time for the update can be completed, and its actions. • Contact our service organization at 1-800-888-7436 if you have any questions, and its actions. • Review, understand, and confirm acknowledgement of the information and instructions outlined in this Urgent Safety Advisory Notice, and its actions.
Siemens | Siemens Medical Solutions USA, Inc | software syngo.CT VB20 in the installed base, with or without syngo.CT software versions VB20_SP1, SP2, SP3, or SP4, may result in workflow interruptions***resulting in the delay in diagnosis, the necessity for patient rescans, and/or the need for additional contrast media may occur | Action Siemens Medical Solutions USA, Inc. initiated a Customer Safety Advisory Notice (CSAN) to customers via CT021/21/S on 8/30/2021. The CSAN informs the customer of the issues health risk, the new software update version VB20_SP5 to eliminate the root causes of these problems. The software update will be distributed free of charge via CT020/21/S and CT071/21/S (update is dependent on the current software version) to all affected customers. ***Action to take:***Siemens Healthineers has developed software update syngo.CT VB20 SP5. This new software contains workflow improvements and solutions to the previously described software issues. Please ask your local service organization if you are interested in more detailed information about the content of the software update.***The new software, VB20_SP5, will be provided free of charge and will be distributed via Update CT020/21/S or CT071/20/S. Update CT071/20/S is only necessary if the system is being updated from software version VB20, VB20_SP1, or VB20_SP2 directly to VB20_SP5. Following the corrective action process CAPA, the cause of identified***product issues has been eliminated and a recurrence of the identified issues is prevented.***This update will be performed remotely. The remote software updates process will require approximately 45 minutes for completion and is completely automatic; however, please make sure the system and power are stable before and during the process, please do not switch off the system during the update process. Siemens highly recommends***starting the installation when the scanner will not be in use or when the necessary time for the update to be completed can be scheduled.***If you have any questions, please contact our service organization at at 1-800-888-7436.******Please review, understand, and confirm acknowledgement of the information and instructions outlined in this***Urgent Safety Advisory Notice. By signing this document, you are acknowledging that you have read and understand the content therein | SoftwareDesign | • Delay in diagnosis actions may occur due to workflow interruptions. • Necessity for patient rescans actions may be required due to software malfunction. • Need for additional contrast media actions may arise from system failure. | • Ask your local service organization for more information about the content of the software update syngo.CT VB20 SP5, which contains workflow improvements and solutions to previously described software issues, as part of the actions. • Ensure the system and power are stable before and during the remote software update process, which will take approximately 45 minutes to complete, as part of the actions. • Schedule the installation when the scanner will not be in use to allow for the necessary time for the update to be completed, as part of the actions. • Contact our service organization at 1-800-888-7436 if you have any questions, as part of the actions. • Review, understand, and confirm acknowledgement of the information and instructions outlined in this Urgent Safety Advisory Notice, as part of the actions.
Siemens | Siemens Medical Solutions USA, Inc | software syngo.CT VB20 in the installed base, with or without syngo.CT software versions VB20_SP1, SP2, SP3, or SP4, may result in workflow interruptions***resulting in the delay in diagnosis, the necessity for patient rescans, and/or the need for additional contrast media may occur | Action Siemens Medical Solutions USA, Inc. initiated a Customer Safety Advisory Notice (CSAN) to customers via CT021/21/S on 8/30/2021. The CSAN informs the customer of the issues health risk, the new software update version VB20_SP5 to eliminate the root causes of these problems. The software update will be distributed free of charge via CT020/21/S and CT071/21/S (update is dependent on the current software version) to all affected customers. ***Action to take:***Siemens Healthineers has developed software update syngo.CT VB20 SP5. This new software contains workflow improvements and solutions to the previously described software issues. Please ask your local service organization if you are interested in more detailed information about the content of the software update.***The new software, VB20_SP5, will be provided free of charge and will be distributed via Update CT020/21/S or CT071/20/S. Update CT071/20/S is only necessary if the system is being updated from software version VB20, VB20_SP1, or VB20_SP2 directly to VB20_SP5. Following the corrective action process CAPA, the cause of identified***product issues has been eliminated and a recurrence of the identified issues is prevented.***This update will be performed remotely. The remote software updates process will require approximately 45 minutes for completion and is completely automatic; however, please make sure the system and power are stable before and during the process, please do not switch off the system during the update process. Siemens highly recommends***starting the installation when the scanner will not be in use or when the necessary time for the update to be completed can be scheduled.***If you have any questions, please contact our service organization at at 1-800-888-7436.******Please review, understand, and confirm acknowledgement of the information and instructions outlined in this***Urgent Safety Advisory Notice. By signing this document, you are acknowledging that you have read and understand the content therein | SoftwareDesign | • Workflow interruptions resulting in delayed diagnosis actions • Necessity for patient rescans actions • Need for additional contrast media actions | • To eliminate health risks, take the following actions: ask your local service organization about the software update syngo.CT VB20 SP5. • Take the following actions: please make sure the system and power are stable before and during the remote software update process. • Take the following actions: do not switch off the system during the update process, and schedule the update when the scanner will not be in use. • If you have any questions, take the following actions: contact our service organization at 1-800-888-7436. • Take the following actions: review, understand, and confirm acknowledgement of the information and instructions outlined in this Urgent Safety Advisory Notice.
Siemens | Siemens Medical Solutions USA, Inc | software syngo.CT VB20 in the installed base, with or without syngo.CT software versions VB20_SP1, SP2, SP3, or SP4, may result in workflow interruptions***resulting in the delay in diagnosis, the necessity for patient rescans, and/or the need for additional contrast media may occur | Action Siemens Medical Solutions USA, Inc. initiated a Customer Safety Advisory Notice (CSAN) to customers via CT021/21/S on 8/30/2021. The CSAN informs the customer of the issues health risk, the new software update version VB20_SP5 to eliminate the root causes of these problems. The software update will be distributed free of charge via CT020/21/S and CT071/21/S (update is dependent on the current software version) to all affected customers. ***Action to take:***Siemens Healthineers has developed software update syngo.CT VB20 SP5. This new software contains workflow improvements and solutions to the previously described software issues. Please ask your local service organization if you are interested in more detailed information about the content of the software update.***The new software, VB20_SP5, will be provided free of charge and will be distributed via Update CT020/21/S or CT071/20/S. Update CT071/20/S is only necessary if the system is being updated from software version VB20, VB20_SP1, or VB20_SP2 directly to VB20_SP5. Following the corrective action process CAPA, the cause of identified***product issues has been eliminated and a recurrence of the identified issues is prevented.***This update will be performed remotely. The remote software updates process will require approximately 45 minutes for completion and is completely automatic; however, please make sure the system and power are stable before and during the process, please do not switch off the system during the update process. Siemens highly recommends***starting the installation when the scanner will not be in use or when the necessary time for the update to be completed can be scheduled.***If you have any questions, please contact our service organization at at 1-800-888-7436.******Please review, understand, and confirm acknowledgement of the information and instructions outlined in this***Urgent Safety Advisory Notice. By signing this document, you are acknowledging that you have read and understand the content therein | SoftwareDesign | • Workflow interruptions may occur, leading to delayed actions. • Delay in diagnosis may result, requiring additional actions. • Patient rescans may be necessary, necessitating repeated actions. • The need for additional contrast media may occur, requiring extra actions. | • Please ask your local service organization for more detailed information about the content of the software update, and to understand the actions you need to take. • Contact our service organization at 1-800-888-7436 if you have any questions, and to discuss the actions required. • Schedule the software update installation when the scanner will not be in use, and ensure the actions outlined are followed. • Make sure the system and power are stable before and during the update process, and take the actions necessary to prevent interruption. • Confirm acknowledgement of the information and instructions outlined in this Urgent Safety Advisory Notice, and take the actions required.
Siemens | Siemens Medical Solutions USA, Inc | software syngo.CT VB20 in the installed base, with or without syngo.CT software versions VB20_SP1, SP2, SP3, or SP4, may result in workflow interruptions***resulting in the delay in diagnosis, the necessity for patient rescans, and/or the need for additional contrast media may occur | Action Siemens Medical Solutions USA, Inc. initiated a Customer Safety Advisory Notice (CSAN) to customers via CT021/21/S on 8/30/2021. The CSAN informs the customer of the issues health risk, the new software update version VB20_SP5 to eliminate the root causes of these problems. The software update will be distributed free of charge via CT020/21/S and CT071/21/S (update is dependent on the current software version) to all affected customers. ***Action to take:***Siemens Healthineers has developed software update syngo.CT VB20 SP5. This new software contains workflow improvements and solutions to the previously described software issues. Please ask your local service organization if you are interested in more detailed information about the content of the software update.***The new software, VB20_SP5, will be provided free of charge and will be distributed via Update CT020/21/S or CT071/20/S. Update CT071/20/S is only necessary if the system is being updated from software version VB20, VB20_SP1, or VB20_SP2 directly to VB20_SP5. Following the corrective action process CAPA, the cause of identified***product issues has been eliminated and a recurrence of the identified issues is prevented.***This update will be performed remotely. The remote software updates process will require approximately 45 minutes for completion and is completely automatic; however, please make sure the system and power are stable before and during the process, please do not switch off the system during the update process. Siemens highly recommends***starting the installation when the scanner will not be in use or when the necessary time for the update to be completed can be scheduled.***If you have any questions, please contact our service organization at at 1-800-888-7436.******Please review, understand, and confirm acknowledgement of the information and instructions outlined in this***Urgent Safety Advisory Notice. By signing this document, you are acknowledging that you have read and understand the content therein | SoftwareDesign | • Delayed diagnosis actions due to workflow interruptions. • Necessity for patient rescans actions resulting from software malfunction. • Need for additional contrast media actions caused by faulty software operation. | • Ask your local service organization for more detailed information about the software update syngo.CT VB20 SP5, which contains workflow improvements and solutions to the previously described software issues as part of the actions. • Ensure the system and power are stable before and during the remote software update process, which will take approximately 45 minutes to complete, as part of the actions. • Start the installation when the scanner will not be in use or when the necessary time for the update to be completed can be scheduled as part of the actions. • Contact our service organization at 1-800-888-7436 if you have any questions as part of the actions. • Review, understand, and confirm acknowledgement of the information and instructions outlined in this Urgent Safety Advisory Notice as part of the actions.
Siemens | Siemens Medical Solutions USA, Inc | software syngo.CT VB20 in the installed base, with or without syngo.CT software versions VB20_SP1, SP2, SP3, or SP4, may result in workflow interruptions***resulting in the delay in diagnosis, the necessity for patient rescans, and/or the need for additional contrast media may occur | Action Siemens Medical Solutions USA, Inc. initiated a Customer Safety Advisory Notice (CSAN) to customers via CT021/21/S on 8/30/2021. The CSAN informs the customer of the issues health risk, the new software update version VB20_SP5 to eliminate the root causes of these problems. The software update will be distributed free of charge via CT020/21/S and CT071/21/S (update is dependent on the current software version) to all affected customers. ***Action to take:***Siemens Healthineers has developed software update syngo.CT VB20 SP5. This new software contains workflow improvements and solutions to the previously described software issues. Please ask your local service organization if you are interested in more detailed information about the content of the software update.***The new software, VB20_SP5, will be provided free of charge and will be distributed via Update CT020/21/S or CT071/20/S. Update CT071/20/S is only necessary if the system is being updated from software version VB20, VB20_SP1, or VB20_SP2 directly to VB20_SP5. Following the corrective action process CAPA, the cause of identified***product issues has been eliminated and a recurrence of the identified issues is prevented.***This update will be performed remotely. The remote software updates process will require approximately 45 minutes for completion and is completely automatic; however, please make sure the system and power are stable before and during the process, please do not switch off the system during the update process. Siemens highly recommends***starting the installation when the scanner will not be in use or when the necessary time for the update to be completed can be scheduled.***If you have any questions, please contact our service organization at at 1-800-888-7436.******Please review, understand, and confirm acknowledgement of the information and instructions outlined in this***Urgent Safety Advisory Notice. By signing this document, you are acknowledging that you have read and understand the content therein | SoftwareDesign | • Workflow interruptions actions may result in delay in diagnosis. • Delay in diagnosis actions may lead to the necessity for patient rescans. • Necessity for patient rescans actions may result in the need for additional contrast media. | • Ask your local service organization for more information about the software update syngo.CT VB20 SP5, which contains workflow improvements and solutions to previously described software issues, as part of the actions. • Ensure the system and power are stable before and during the remote software update process, which will require approximately 45 minutes for completion, as part of the actions. • Schedule the installation when the scanner will not be in use or when the necessary time for the update to be completed can be scheduled, as part of the actions. • Do not switch off the system during the update process, as part of the actions. • Contact our service organization at 1-800-888-7436 if you have any questions, as part of the actions. • Review, understand, and confirm acknowledgement of the information and instructions outlined in this Urgent Safety Advisory Notice, as part of the actions.
Oncentra Brachy | Elekta, Inc. | Under certain circumstances, there can be a difference between the planned and delivered dose of the two ovoid channels. | Action An Urgent Important Field Safety Notification letter (FAC-NU-0007) was issued to consignees to notify of the issue, clinical impact and recommend the following:******1) Recommended User Action in the unique scenario, which is only applicable for the Venezia" applicator: Avoid use of the combination of the two Venezia" applicator Anchor Points O1-Center and O2-Center by disabling one of them in the Applicator Library Manager.******2)Recommended User Action in the exceptional situation that may also occur for other Applicator Models: Avoid incorrect placement of the Anchor Points, especially exchanging left / right anchor point positions.******3) Recommended User Actions for both scenarios:***Until an updated version of Oncentra Brachy is available, it is strongly advised:***" To use additional Anchor Points during applicator modelling. This will further reduce the chance of***mirroring the Applicator Model.***" To cross-check the source dwell position information from the Case Explorer to be in accordance***with the source dwell positions in the image views.***" Use the Triangles option for the applicator skin visualization in the 3D view, and review whether the***Applicator Model has a correct shading display.***" Perform proper Quality Assurance for all treatment plans before treatment delivery.*********Consignees are required to acknowledge receipt of the notification to Elekta within 30 days. | SoftwareDesign | • Difference between planned and delivered dose of the two ovoid channels may occur, affecting treatment actions. • Delivered dose may not match the intended dose, potentially impacting patient safety and treatment outcomes actions. | • Recommended actions to avoid incorrect use of Venezia applicator: disable one Anchor Point (O1-Center or O2-Center) in the Applicator Library Manager. • Recommended actions to avoid incorrect placement of Anchor Points: avoid exchanging left/right anchor point positions. • Recommended actions for applicator modelling: use additional Anchor Points to reduce mirroring of the Applicator Model. • Recommended actions for quality assurance: cross-check source dwell position information from the Case Explorer with image views. • Recommended actions for applicator skin visualization: use the Triangles option in the 3D view and review correct shading display. • Recommended actions before treatment delivery: perform proper Quality Assurance for all treatment plans.
MOSAIQ Oncology Information System | Elekta, Inc. | A drug strength in MOSAIQ can be changed during the ordering process and approved, but the saved drug strength reverts to the original value. | Action The recalling firm issued an "URGENT IMPORTANT FIELD SAFETY NOTIFICATION" dated August 2021 via email to its consignees on 9/2/2021. Recommendations were made to the user until a fix is available. | SoftwareDesign | • Inconsistent drug strength value actions lead to incorrect information retention. • Approved changes to drug strength during ordering process do not persist, causing reversion to original value actions. | • Take actions to follow the recommendations made by the recalling firm until a fix is available.
Siemens | Siemens Medical Solutions USA, Inc | Software versions may result in sporadic problems causing scanning workflow interruptions and unexpected user notifications. Sporadic software errors may also occur during interventional workflows, resulting in delay in diagnosis or scan aborts with the necessity for patient rescan may occur | Action Siemens Medical Solutions USA, Inc. initiated a Customer Safety Advisory Notice (CSAN) to US customers via CT024/21/S on 8/18/2021. The letter states reason for recall, helath risk and action to take:***There are no workarounds available for the identified issues.******Siemens Healthineers has developed software update syngo.CT VA30A_SP4 to ensure uninterrupted scanning workflows and to reduce the number of user notifications. This update will also provide workflow improvements, bug fixes for the performance and stability problems observed in the installed base.***The corrective action will be provided free of charge and will be distributed via of the update packages described above, depending on your current software version.***Following the corrective action, the cause has been eliminated and recurrence of the identified issues are prevented. This software update will be provided to you free of charge.******Software updates CT022/21/S and CT025/21/S will be performed onsite. Updates CT016/21/S and CT018/21/S will be performed remotely. The software updates process will require approximately 180 minutes for completion.***Please make sure the system and power are stable before and during the process. Please do not switch off the system during the update process. Siemens highly recommends starting the installation when the scanner is not in use or when the necessary time for the update to be completed can be scheduled. | SoftwareDesign | • Scanning workflow interruptions and unexpected user notifications may occur due to software version issues, affecting scanning actions. • Sporadic software errors during interventional workflows may result in delayed diagnosis or scan aborts, necessitating patient rescan actions. • Software errors may lead to patient rescan actions, potentially causing inconvenience and discomfort to patients. | • To ensure uninterrupted scanning workflows and reduce user notifications, Siemens Healthineers has developed software update syngo.CT VA30A_SP4, which provides workflow improvements, bug fixes, and eliminates the cause of identified issues; take actions to install this update. • The corrective actions will be provided free of charge and will be distributed via update packages, take actions to receive and install them. • Take actions to ensure the system and power are stable before and during the update process; do not switch off the system during the update process. • Schedule the update when the scanner is not in use or when the necessary time for the update to be completed can be allocated; take actions to complete the update process.
Philips | Philips North America Llc | When setting patient weight unit preferences to pounds, SynchRight P3T (Personalized Patient Protocol) software will result in increased contrast volume, resulting in increased contrast volume recommended for the patient. | Action Philips issued Urgent Medical Device Correction letter dated 6/8/21 to consignees via certified mail. Letter states reason for recall, health risk and action to take:***following steps:***Go to Preferences> Regional Settings > Weight units, select the Weight units to Kg.***Step 1: Verify that the patient weight unit is set to Kg, rather than Lbs.***Step 2: Change the weight unit to Kg if weight unit is set to Lbs, referring to the screenshot below***Step 3: Log out and then log back in to make the changes effective.******Please share this notice with all those who need to be aware within your organization.***Please complete and return the attached response form to Philips Healthcare promptly upon receipt and no later than 30 days from receipt, via email to CTNM.QARA@PHILIPS.COM******The software update is planned to be released in the 3rd quarter of 2021. Once the software update is released, a Philips representative will contact you to schedule a software update to permanently resolve the issue.***For North America, contact the Customer Care Solutions Center (1-800-722-9377). | SoftwareDesign | • Incorrect patient weight unit preferences leading to increased contrast volume recommended for patients, potentially causing unsafe actions. • Failure to accurately calculate contrast volume, resulting in potentially harmful actions to patients. | • Take the following actions: Go to Preferences > Regional Settings > Weight units, select the Weight units to Kg. • Take the actions: Verify that the patient weight unit is set to Kg, rather than Lbs. • Perform the actions: Change the weight unit to Kg if weight unit is set to Lbs. • Take the actions: Log out and then log back in to make the changes effective. • Share this notice with all those who need to be aware within your organization and complete the attached response form. • Take the actions: Contact the Customer Care Solutions Center (1-800-722-9377) for assistance.
Philips | Philips North America Llc | When setting patient weight unit preferences to pounds, SynchRight P3T (Personalized Patient Protocol) software will result in increased contrast volume, resulting in increased contrast volume recommended for the patient. | Action Philips issued Urgent Medical Device Correction letter dated 6/8/21 to consignees via certified mail. Letter states reason for recall, health risk and action to take:***following steps:***Go to Preferences> Regional Settings > Weight units, select the Weight units to Kg.***Step 1: Verify that the patient weight unit is set to Kg, rather than Lbs.***Step 2: Change the weight unit to Kg if weight unit is set to Lbs, referring to the screenshot below***Step 3: Log out and then log back in to make the changes effective.******Please share this notice with all those who need to be aware within your organization.***Please complete and return the attached response form to Philips Healthcare promptly upon receipt and no later than 30 days from receipt, via email to CTNM.QARA@PHILIPS.COM******The software update is planned to be released in the 3rd quarter of 2021. Once the software update is released, a Philips representative will contact you to schedule a software update to permanently resolve the issue.***For North America, contact the Customer Care Solutions Center (1-800-722-9377). | SoftwareDesign | • Inaccurate contrast volume recommendations due to incorrect weight unit conversions lead to improper patient treatment actions. • Increased contrast volume recommended for patients results in potential harm or adverse reactions to patients actions. | • Take the following actions: Go to Preferences > Regional Settings > Weight units, select the Weight units to Kg, and verify that the patient weight unit is set to Kg. • Take actions to change the weight unit to Kg if it is set to Lbs, referring to the provided screenshot. • Take actions to log out and then log back in to make the changes effective. • Share this notice with all those who need to be aware within your organization and take actions to complete and return the attached response form promptly. • Take actions to contact Philips Healthcare via email to CTNM.QARA@PHILIPS.COM for further assistance. • Take actions to contact the Customer Care Solutions Center (1-800-722-9377) for support.
Philips | Philips North America Llc | When setting patient weight unit preferences to pounds, SynchRight P3T (Personalized Patient Protocol) software will result in increased contrast volume, resulting in increased contrast volume recommended for the patient. | Action Philips issued Urgent Medical Device Correction letter dated 6/8/21 to consignees via certified mail. Letter states reason for recall, health risk and action to take:***following steps:***Go to Preferences> Regional Settings > Weight units, select the Weight units to Kg.***Step 1: Verify that the patient weight unit is set to Kg, rather than Lbs.***Step 2: Change the weight unit to Kg if weight unit is set to Lbs, referring to the screenshot below***Step 3: Log out and then log back in to make the changes effective.******Please share this notice with all those who need to be aware within your organization.***Please complete and return the attached response form to Philips Healthcare promptly upon receipt and no later than 30 days from receipt, via email to CTNM.QARA@PHILIPS.COM******The software update is planned to be released in the 3rd quarter of 2021. Once the software update is released, a Philips representative will contact you to schedule a software update to permanently resolve the issue.***For North America, contact the Customer Care Solutions Center (1-800-722-9377). | SoftwareDesign | • Incorrect patient weight unit preferences in SynchRight P3T software lead to actions resulting in increased contrast volume recommended for patients. • Increased contrast volume due to software malfunction affects the system's operation and patient safety through unnecessary or excessive contrast administration actions. | • Take the following actions: Go to Preferences > Regional Settings > Weight units, select the Weight units to Kg. • Take the actions to Verify that the patient weight unit is set to Kg, rather than Lbs. • Take the actions to Change the weight unit to Kg if weight unit is set to Lbs, referring to the screenshot below. • Take the actions to Log out and then log back in to make the changes effective. • Take the actions to Share this notice with all those who need to be aware within your organization. • Take the actions to Complete and return the attached response form to Philips Healthcare promptly upon receipt. • Take the actions to Contact the Customer Care Solutions Center for further assistance.
Philips | Philips North America Llc | When setting patient weight unit preferences to pounds, SynchRight P3T (Personalized Patient Protocol) software will result in increased contrast volume, resulting in increased contrast volume recommended for the patient. | Action Philips issued Urgent Medical Device Correction letter dated 6/8/21 to consignees via certified mail. Letter states reason for recall, health risk and action to take:***following steps:***Go to Preferences> Regional Settings > Weight units, select the Weight units to Kg.***Step 1: Verify that the patient weight unit is set to Kg, rather than Lbs.***Step 2: Change the weight unit to Kg if weight unit is set to Lbs, referring to the screenshot below***Step 3: Log out and then log back in to make the changes effective.******Please share this notice with all those who need to be aware within your organization.***Please complete and return the attached response form to Philips Healthcare promptly upon receipt and no later than 30 days from receipt, via email to CTNM.QARA@PHILIPS.COM******The software update is planned to be released in the 3rd quarter of 2021. Once the software update is released, a Philips representative will contact you to schedule a software update to permanently resolve the issue.***For North America, contact the Customer Care Solutions Center (1-800-722-9377). | SoftwareDesign | • Inaccurate contrast volume calculation actions. • Increased recommended contrast volume actions. • Potential harm to patients due to excessive contrast administration actions. | • Take the following actions: Go to Preferences > Regional Settings > Weight units, and select Kg as the weight unit. • Take the actions of verifying that the patient weight unit is set to Kg, rather than Lbs. • Take the actions of changing the weight unit to Kg if it is currently set to Lbs, referring to the provided screenshot. • Take the actions of logging out and then logging back in to make the changes effective. • Take the actions of sharing this notice with all those who need to be aware within your organization. • Take the actions of completing and returning the attached response form to Philips Healthcare promptly upon receipt. • Take the actions of contacting the Customer Care Solutions Center (1-800-722-9377) for North America.
Siemens | Siemens Medical Solutions USA, Inc | Software error VE10B -In case of a fault within the drivetrain of an axis and while performing a stepped peripheral imaging (PERISTEPPING/PERIVISION), the C-arm may leave its intended travel path by up to 22 cm (9") result in canceling clinical treatment or to continue treatment on an alternative system | Action Siemens issued a Urgent Medical Device Correction Letter on 5/27/21 via AX064/20/S. Additionally, a software solution is being distributed via Update Instruction AX063/20/S to all affected customers. Letter states reason for recall, health risk and action to take:******Planned procedures may have to be terminated and performed on an alternative X-ray system. Please follow the standard emergency procedures you have in place in the event of system failures. It is recommended to consider reinforcing these emergency procedures until the update has been completed.************Siemens will update the affected software to prevent this non-conformity from reoccurring via Update Instruction AX063/20/S. When the system recognizes a deviation of the intended travel path, unit movements will be stopped immediately, similar to an activation of a proximity switch.******Our service organization will contact you shortly to arrange a date to perform this corrective action. Please feel free to contact our service organization for an earlier appointment at 1-800-888-7436. Following the correction, the cause is eliminated, and any recurrence of this potential fault is prevented.******Acknowledge Receipt of this Safety Advisory Notice******Please review, understand, and confirm acknowledgement of the information and instructions outlined in this Urgent Safety Advisory Notice. By signing this document, you are acknowledging that you have read and understand the content therein.******Request that you promptly notify and instruct accordingly all the staff at your organization who need to be aware of this notice and will comply with the recommendations therein. Please ensure that this safety advisory is retained in your product related records appropriately. In addition, if you may have further distributed this product, please identify your customers, and notify them at once of this product recall. If the device has been sold and is therefore no longer in your possession, please forward this notice to the new owner. | SoftwareDesign | • C-arm may deviate from its intended travel path by up to 22 cm during stepped peripheral imaging actions. • Fault within the drivetrain of an axis can cause the C-arm to leave its intended path actions. • The deviation can result in canceling clinical treatment actions. • Alternative treatment systems may be required actions. | • Take actions to terminate planned procedures and perform them on an alternative X-ray system if necessary, following standard emergency procedures. • Consider reinforcing emergency procedures until the software update is completed, taking necessary actions to ensure safety. • Arrange for the software update with our service organization, taking actions to prevent the recurrence of the potential fault. • Review, understand, and acknowledge receipt of this Safety Advisory Notice, taking actions to comply with the instructions. • Notify and instruct all staff who need to be aware of this notice, taking actions to ensure they comply with the recommendations. • Retain this safety advisory in your product-related records, taking actions to ensure accessibility. • If you have further distributed this product, identify your customers and notify them of the product recall, taking actions to inform them of the necessary actions.
Change Healthcare Enterprise Viewer | CHANGE HEALTHCARE CANADA COMPANY | A software defect was identified where the Image Styles defined by presentation states from CHRS are not displayed in CHEV. | Action On 04/19/21 the firm sent a letter to its consignees with the following instructions:************Immediate actions to take to reduce the risk to patient******Prior to treating patients, always confirm the correct patient anatomy by reviewing the laterality markers applied during image acquisition.******When the issue is detected, the clinician is able to manually adjust the Image Styles in CHEV to accommodate their viewing requirements and/or orient images to standard anatomical position.************Product updates that will address the issue******Change Healthcare is developing a software update to correct image display of orientation, image inversion and window level information from Image Styles.******Recommended actions to take on potentially affected data******We strongly encourage customers receiving this Urgent Field Safety Notice to immediately contact Change Healthcare Customer Support to:******1) Acknowledge they have read and understood this Urgent Field Safety Notice; and,******2) Arrange the installation of software update.******Change Healthcare Customer Support: 1-800-663-2533 Menu Option 3 (CHEV) | SoftwareDesign | • Failure to display Image Styles defined by presentation states from CHRS in CHEV affects system operation by preventing users from accessing expected visual information, requiring corrective actions. • Inability to display Image Styles results in inconsistent presentation, potentially leading to user confusion or misinformation, necessitating immediate corrective actions. • Defect in Image Style display compromises system performance, demanding prompt investigation and corrective actions. | • Take immediate actions to reduce the risk to patient by confirming the correct patient anatomy and manually adjusting Image Styles in CHEV. • Take actions to accommodate viewing requirements and/or orient images to standard anatomical position when the issue is detected. • Take actions to contact Change Healthcare Customer Support to acknowledge receipt of this Urgent Field Safety Notice and arrange the installation of software update. • Take actions to call Change Healthcare Customer Support at 1-800-663-2533 Menu Option 3 (CHEV) for assistance.
RaySearch | RAYSEARCH LABORATORIES AB | For some LINAC types, merging clinical beams with beams of approximate dose may lead to the approximate dose erroneously labeled as clinical dose. Merge beams can be used manually, in scripting, or as part of the***Automated breast planning feature. In some cases, when merging two beams where the first beam has clinical dose and the second beam has approximate dose, the dose of the resulting beam will be labeled as Clinical, although dose for some of the beam segments is still calculated with the SVD dose engine and it should be labeled Approximate: Mixed dose. The difference between approximate and final Clinical dose is in most cases small, but there can be body sites such as lung where the difference can be significant. | Action RaySearch issued Field Safety Notice, Medical Device Correction #74977 letter on 22 February 2021 via Email. Letter states reason for recall, health risk and action to take:******ACTIONS TO BE TAKEN BY THE USER******" Manual ***Merge beams*** in the 3D-CRT module Scenario: When merging a beam B1 with other beams (e.g., B2 and B3) in the 3D-CRT module.******Action: If the target beam B1 has a clinical dose before the merge, make sure that all the beamsto be merged (B2 and B3) also have a clinical dose computed with the same dose engine.******" Scripting when using the action MergeBeamSegments******Scenario: When running a script containing the action MergeBeamSegments as below:******beam_set.MergeBeamSegments(TargetBeamName="B1", MergeBeamNames=["B2", "B3"])******Action: If the target beam B1 has a clinical dose before running the script, make sure that the beams to be merged (B2 and B3) also have a clinical dose computed with the same dose engine.******" Automated breast planning******Scenario: When creating a treatment plan in the Automated breast planning module, the generated breast plan has two beams: A1 and A2.******Actions:******1. Change module to Plan optimization.******2. Invalidate the dose for both the A1 and A2 beams by changing a beam property that can easily be restored (e.g., change the couch angle from 0 to 1).******3. Set the beam property back to the intended value (e.g., set the couch angle back to 0 for both beams).******4. Compute final dose.******" Educate planning staff and all users about this workaround.******" Inspect your product and identify all installed units with the above software version number(s).******" Confirm that you have read and understood this notice by replying to the notification email.******SOLUTION******This issue will be resolved in the next version of RayStation/RayPlan, scheduled for market release in May 2021 (subject to market clearance in some markets). If customers wish to continue using versions of******RayStation/RayPlan affected by this notice, all users must maintain awareness of this notice. | SoftwareDesign | • Inaccurate labeling of dose type occurs when merging clinical beams with beams of approximate dose, leading to potential misuse of clinical dose actions. • Merged beams may be incorrectly labeled as "Clinical" despite containing approximate dose calculations, affecting treatment planning actions. • The difference in dose calculation can be significant in certain body sites, such as lung, potentially impacting patient treatment actions. | • Actions to be taken by the user: When merging a beam B1 with other beams (e.g., B2 and B3) in the 3D-CRT module, actions require that all beams to be merged also have a clinical dose computed with the same dose engine. • Actions when scripting using the action MergeBeamSegments: If the target beam B1 has a clinical dose before running the script, actions require that the beams to be merged also have a clinical dose computed with the same dose engine. • Actions for automated breast planning: Actions require changing the module to Plan optimization, invalidating the dose for both beams, changing a beam property, setting the beam property back, and computing the final dose. • Educate planning staff and all users about these actions and workarounds.
Change Healthcare Enterprise Viewer | CHANGE HEALTHCARE CANADA COMPANY | Change Healthcare has identified an intermittent software defect which may result in an anchor study failing to display. | Action On 3/15/21 the firm sent an Urgent Field Safety Notice to its customers with the following instructions:************Immediate actions to take to reduce the risk to patient******When the issue occurs, the initial anchor images are populated in the viewport. The user can select the series identifier in the viewport to display other series for the study.************If the initial anchor images in the viewport have been replaced with reference images, the user does not have the ability to view the anchor series. However, the user can reopen the study from the EHR or perform a specific search in CHEV to allow the anchor study series to display in the viewport. In addition, the user can change the timeline filter to omit the scheduled study, which would allow the anchor study to display in the timeline.************Product updates that will address the issue******Change Healthcare has developed a software update to correct this issue.************Recommended actions to take on potentially affected data******We strongly encourage customers receiving this Urgent Field Safety Notice to immediately contact Change Healthcare Customer Support to:******1) Acknowledge they have read and understood this Urgent Field Safety Notice; and,******2) Arrange the installation of software update. | SoftwareDesign | • Anchor study failing to display due to intermittent software defect may cause incorrect actions. • Failure to display anchor study may lead to incomplete or inaccurate actions. | • Immediate actions to take to reduce the risk to patient include selecting the series identifier in the viewport to display other series for the study. • Take actions to reopen the study from the EHR or perform a specific search in CHEV to allow the anchor study series to display in the viewport. • Take actions to change the timeline filter to omit the scheduled study, which would allow the anchor study to display in the timeline. • Take actions to contact Change Healthcare Customer Support to acknowledge reading and understanding this Urgent Field Safety Notice and arrange the installation of software update.
Siemens | Siemens Medical Solutions USA, Inc | Due to the configuration of certain Windows Service Permissions within the operating systems of the Sensis/ Sensis Vibe computer, there is a risk for exposure of sensitive information, manipulation of data, or Denial of Service attacks and could result in incorrect diagnostic or therapeutic decisions | Action Siemens Healthineers issued Customer Safety Advisory Notice dated 3/4/21 to affected customers via AX074/20/S. Letter states reason for recall, health risk and action to take:******Siemens will restrict the Windows Service Permissions to the required level only and the IT security vulnerability of your system will be reduced via Update Instruction AX073/20/S.******Our service organization will contact you shortly to arrange a date to perform this corrective action. Please feel free to contact our service organization for an earlier appointment at 1-800-888-7436.************Please review, understand, and confirm acknowledgement of the information and instructions outlined in this Urgent Safety Advisory Notice. By signing this document, you are acknowledging that you have read and understand the content therein.******. | SoftwareDesign | • Risk of exposure of sensitive information through certain Windows Service Permissions actions. • Potential for manipulation of data through Windows Service Permissions actions. • Vulnerability to Denial of Service attacks through Windows Service Permissions actions. • Possibility of incorrect diagnostic or therapeutic decisions due to Windows Service Permissions actions. | • Siemens will restrict the Windows Service Permissions to the required level only to reduce the IT security vulnerability of your system through Update Instruction, and our service organization will contact you to arrange a date to perform these actions. • Please contact our service organization for an earlier appointment at 1-800-888-7436 to schedule these actions. • Review, understand, and confirm acknowledgement of the information and instructions outlined in this Urgent Safety Advisory Notice to proceed with these actions. • Sign this document to acknowledge that you have read and understand the content and will take these actions.
Brainlab | Brainlab AG | The ExacTrac 6.x Patient Positioning System may incorrectly position the patient when using the ExacTrac Cone Beam CT (CBCT) with a TrueBeam-specific optional subvolume-CBCT. | Action BrainLab sent an FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letters dated February 16, 2015, to all affected customers. The letters included instructions for customers to: 1) Do not use any actively re-reconstructed CBCT subvolumes from the Varian TrueBeam system with the Brainlab ExacTrac 6.x CBCT Import & Alignment Software module; 2) Import exclusively original, not modified, CBCT volumes into ExacTrac to use for patient positioning; and, 3) Continue to verify any ExacTrac CBCT-based correction using the ExacTrac X-ray verification and/or an external IGRT procedure as mandatory. ************Brainlab provided existing potentially affected ExacTrac v.6.x CBCT Import & Alignment Software customers (with a Varian TrueBeam system) with the product notification information. Brainlab will provide a software update with this issue solved to affected customers. Brainlab will actively contact the customer starting August 2015 to schedule the update installation.************For further questions please call Customer Hotline + 49 89 99 15 68 44 or 1-800 597-5911 ( For US Customers only) | SoftwareInTheUseEnviorenment | • Inaccurate patient positioning actions occur when using the ExacTrac 6.x Patient Positioning System with the TrueBeam-specific optional subvolume-CBCT feature. • This malfunction can lead to incorrect treatment actions, potentially compromising patient safety and treatment outcomes. | • Take the following actions: Do not use any actively re-reconstructed CBCT subvolumes from the Varian TrueBeam system with the Brainlab ExacTrac 6.x CBCT Import & Alignment Software module. • Take the following actions: Import exclusively original, not modified, CBCT volumes into ExacTrac to use for patient positioning. • Take the following actions: Continue to verify any ExacTrac CBCT-based correction using the ExacTrac X-ray verification and/or an external IGRT procedure as mandatory.
Brainlab | Brainlab AG | iPlan RT Radiation Treatment Planning Software: Potentially incorrect patient positioning when using multiple localized CT image data sets. | Action Brainlab sent a FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letter dated November 19, 2014, to all affected customers. The letter included instructions to: ************1) If clinically not required, do not use multiple localized CT scans within one treatment plan (to avoid potentially incorrect Reference and Alignment Set assignments in the first place). ************2) If you must use multiple localized CT scans, e.g. for recurrent treatment planning of the same patient, always guarantee that the latest (see note *) CT scan is defined as both Alignment Set and Reference Set during treatment planning. Brainlab will provide a software solution to prevent the described scenario from occurring. ************Brainlab will actively contact affected customers tentatively starting January 2016 to schedule the update. Brainlab will additionally refine the instructions for use of the iPlan RT treatment planning software in regard to implications of Alignment Set and Reference Set selections and will provide this user guide update to existing affected customers together with the software update.************Customers with questions can contact Brainlab at 800-597-5911 or via e-mail at us.support@brainlab.com. | SoftwareInTheUseEnviorenment | • Incorrect patient positioning actions due to software issue when using multiple localized CT image data sets. • Potentially compromised treatment outcomes actions resulting from incorrect patient positioning. • Risk of incorrect radiation treatment actions administered to patients due to software malfunction. | • To avoid incorrect Reference and Alignment Set assignments, take the following actions: do not use multiple localized CT scans within one treatment plan if clinically not required. • If multiple localized CT scans are necessary, take the following actions: ensure the latest CT scan is defined as both Alignment Set and Reference Set during treatment planning.
Brainlab | Brainlab AG | iPlan RT Radiation Treatment Planning Software: Potentially incorrect patient positioning when using multiple localized CT image data sets. | Action Brainlab sent a FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letter dated November 19, 2014, to all affected customers. The letter included instructions to: ************1) If clinically not required, do not use multiple localized CT scans within one treatment plan (to avoid potentially incorrect Reference and Alignment Set assignments in the first place). ************2) If you must use multiple localized CT scans, e.g. for recurrent treatment planning of the same patient, always guarantee that the latest (see note *) CT scan is defined as both Alignment Set and Reference Set during treatment planning. Brainlab will provide a software solution to prevent the described scenario from occurring. ************Brainlab will actively contact affected customers tentatively starting January 2016 to schedule the update. Brainlab will additionally refine the instructions for use of the iPlan RT treatment planning software in regard to implications of Alignment Set and Reference Set selections and will provide this user guide update to existing affected customers together with the software update.************Customers with questions can contact Brainlab at 800-597-5911 or via e-mail at us.support@brainlab.com. | SoftwareInTheUseEnviorenment | • Incorrect patient positioning may occur due to inconsistent handling of multiple localized CT image data sets, leading to potentially incorrect treatment actions. • Inaccurate treatment planning may result from the software's inability to correctly integrate multiple localized CT image data sets, affecting treatment actions. • Patients may receive incorrect radiation treatment due to the software's defect, compromising their safety and treatment outcomes through ineffective actions. | • Take actions to avoid using multiple localized CT scans within one treatment plan if clinically not required. • Take actions to guarantee that the latest CT scan is defined as both Alignment Set and Reference Set during treatment planning if you must use multiple localized CT scans. • Contact Brainlab at 800-597-5911 or via e-mail at us.support@brainlab.com for questions and actions.
Brainlab | Brainlab AG | Brainlab Navigation Software Spine &amp; Trauma 3D 2.0/2.1 offers automatic registration of intraoperatively acquired CT image data sets. To enable automatic registration, the software requires the gantry position of the scanner. ******The gantry position can either be entered manually or submitted automatically from the CT scanner. If automatic gantry communication is available, the so-called &quot | Action The firm, Brainlab AG, sent "FIELD SAFETY NOTICE / PRODUCT NOTIFICATION" letters dated May 21, 2014 to all consignees. The letters included instructions for consignees to: 1) not use Automatic Registration of iCT Data in combination with Siemens SOMATOM Definition AS CT scan (with Software Somaris SOM 7 or higher and Automatic Gantry Communication (available for Spine & Trauma 3D 2.0 and 2.1 only); 2) always make sure to perform careful accuracy verification according to the instructions in the Software User Guide; 3) Brainlab will perform corresponding service visits to change the gantry communication from automatic to manual to enable a safe and effective usage of this device combination, and 4) advise the appropriate personnel working in your department of the content of the letter.************Customers with any questions can contact Brainlab via telephone at 800-597-5911 or e-mail at us.support@brainlab.com. | SoftwareInTheUseEnviorenment | • Inadequate gantry position submission actions may hinder automatic registration of intraoperatively acquired CT image data sets. • Manual entry of gantry position may be required if automatic gantry communication is not available, which could lead to errors or delays in the registration process actions. | • Take the following actions: do not use Automatic Registration of iCT Data in combination with Siemens SOMATOM Definition AS CT scan. • Take the following actions: always make sure to perform careful accuracy verification according to the instructions in the Software User Guide. • Take the following actions: await Brainlab's corresponding service visits to change the gantry communication from automatic to manual to enable a safe and effective usage of this device combination. • Take the following actions: advise the appropriate personnel working in your department of the content of this notification.